## Slide 1: Update on SBS Back Tracker GEM @ UVa
K. Gnanvo, N. Liyanage, V Nelyubin, K. Saenboonruang, Seth Saher
Cross Talk issues with APV25 Electronics
Uniformity tests results with 90Sr source

## Slide 2: SBS Back Tracker Proto1 with APV25-SRS FE cards
Pedestal RMS noise distribution over 16 APV
ADC Ch counts
For 50 × 50 cm2 COMPASS-like readout, typical rms after common mode correction of the baseline is on average of 6-7 adc counts for apv25-SRS
  @ 230 e-/adc  ~ 1200 to 1500 ENC
  a cut at 5 sigma for zero suppression  ~ 6000 e-

## Slide 3: Gain (non) uniformity of the chamber
Y
X
window
Drift cathode
GEM foil 1
GEM foil 2
GEM foil 3
Mean cluster adc counts 2D map
ADC counts
Gas flow uniform in Y direction  gas inlet and outlet
We suspects a non uniformity of the gain due to the gas flow in the X direction
The gas flow in X is facilitated by groove in the spacer but not sure if it is
Y
X

## Slide 4: Cross Talk in APV25 Electronics
Current version of the analysis (AMORE) code:
Only single cluster event used for 2D hit maps, X-Y charges correlation etc … and minimal cluster size = 2hits
For high amplitude hits, cross talk cluster survives a 5s zero suppression cut  Event seen with 2 clusters and rejected
Cross talk is suppressed at 7s here in this example
5s cut
7s cut

## Slide 5: Cross talk effect previously reported at RD51 Coll. meetings 2011
K. Gnanvo, GEM R&D Muon Tomography, Florida Tech
RD51 Coll. Meeting April 2011
M. Byzsewsky, Micromegas for ATLAS Muon chamber Upgrade
RD51 mini week January 201111

## Slide 6: Distribution of the relative position (Xcluster1 – Xcluster2) of two clusters in 2-clusters events
Along X-axis
Along Y-axis
Pics at 32, 88 and 120 
cross talk effect
Real 2-cluster events
Along X-axis
Along Y-axis
Distribution of ADC counts ratio (ADCcluster1 / ADC cluster2 ) of two clusters in 2-clusters events
Large ratio > 20 correlated to the cross talk pics

## Slide 7: 2D Map of “mean” ADC: Gain uniformity
@ 5s pedestal cut
@ 10s pedestal cut
Along X-axis
Along Y-axis
ADC(X) / ADC (Y)
= 1.34
Cross talk effect more pronounced for X than for Y  non equal charge sharing  Cross talk cluster suppressed at 5s cut
“mean” ADC
=
ƩN ADC / NHits

## Slide 8: Cross talk only for APV25-slaves  Effect of SAMTEC flat cable?
M
S
M
S
SAMTEC flat cable

## Slide 9: Test with 90Sr source
Hit map (NHits / bin)
“mean” ADC (gain uniformity)
Chart: Efficiency
| 3950 | 4000 | 4050 | 4100 | 4150 | 4200 | 4250 | 4300 | 4350 | 4400
A11 | 0.42329 | 0.54962 | 0.66043 | 0.74907 | 0.83218 | 0.87207 | 0.88574 | 0.90421 | 0.9345 | 0.93523
A2 | 0.46466 | 0.61352 | 0.72802 | 0.79561 | 0.85213 | 0.88832 | 0.92194 | 0.93043 | 0.94004 | 0.9537
B11 | 0.41454 | 0.53333 | 0.66475 | 0.77323 | 0.82032 | 0.8621 | 0.88116 | 0.89634 | 0.91141 | 0.91883
“mean” ADC
=
ƩN ADC / NHits
Efficiency measurement with 90Sr
About 92% at 4400 V
90Sr Source just on top of the spacer along Y axis  few % drop in the efficiency of the chamber expected
Results consistent with the > 95% efficiency expected for the Triple-GEM