## Slide 1: FASTBUS DAQ
Sergey Abrahamyan
University of Virginia
Yerevan Physics Institute
SBS collaboration

## Slide 2: Overview
FASTBUS requirements
Dead time improvements
Hardware preparation and testing
Summary

## Slide 3: experiment
Extrapolation from GeN
Threshold = 1.2 GeV
L(A1N) = 2 × 1036 cm-2 s-1
Trigger rate = 10 kHz

## Slide 4: experiment
Detector | Detector Channels | DAQ channels | Gate [ns] | hits
Gas Cherenkov | 550 | 550 time 550 amplitude | 50 | 30 30
Scintillators | 180 | 180 time 180 amplitude | 50 | 5 5
Calorimeter | 243 | 243 amplitude | 100 | 45
Calorimeter Sums | 26 | 26 amplitude 26 time | 100 | 2 2
120 hits/event = 480 bytes/event at 10kHz trigger rate → 5Mb/s
With TDC full scale window on a level of 1-1.5 µs data rate will be ~50Mb/s
Sparsification mode will allow significantly reduce data rate.
999 amplitude channels - 16 ADC modules or 11 TDC+MQT modules
756 time channels - 8 TDC modules

## Slide 5: experiment
MC simulation
11 GeV 80 µA beam
40 cm LH2
ECal at 30° at 3.0m
32×56 = 1792 blocks
4.0×4.0×40.0 cm
Ethr/Emax (%) | Rate (kHz)
50 | 3500
75 | 320
85 | 120

## Slide 6: experiment
Block threshold (MeV) | Number of fired blocks | Data rate (Mb/s) Ethr/Emax = 0.75 | Data rate (Mb/s) Ethr/Emax = 0.85
0 | 75 | 2×75×4×0.32 = 192 | 2×75×4×0.12 = 72
1 | 53 | 136 | 50
3 | 35 | 90 | 34
5 | 28 | 72 | 26
L1 trigger rate 320 kHz at 75% threshold
120 kHz at 85% threshold
1792 amplitude channels - 28 ADC modules or 19 TDC+MQT modules
1792 time channels - 19 TDC modules

## Slide 7: Methods to reduce dead time
Sparsification
Event blocking
Parallel DAQ systems

## Slide 8: Sparsification
Common Stop
Start from background
Start from signal
full- scale window
0-32 µs (8ns step)
Sparsification
0-8µs
(0.5ns step)
Gate
Typical TDC spectrum without sparsification. Only 300ns window out of 1.5µs full scale window has useful data
Enabling of sparsification allows significantly reduce event size. All TDC modules have been tested to work in sparsification mode.

## Slide 9: Event Blocking
Readout
Triggers
Readout
Overhead
Readout
Overhead
Readout
Overhead
Overhead
Triggers
Readout
Readout
Readout
Readout
Overhead
For one channel readout.
4 events no EB – 280 µs
4 events with EB – 220 µs
20µs 50µs 20µs 50µs 20µs 50µs 20µs 50µs
20µs 50µs 50µs 50µs 50µs
Readout time is 50 µs + 2 µs per 16 channel.
64 channels – 58 µs
128 channels – 66 µs

## Slide 10: Event Blocking test results
Event Blocking | Number of channels | Number of modules | Life time (%)
1 | 1 | 1 | 71
4 | 1 | 1 | 90
1 | 64 | 3 | 65
4 | 64 | 3 | 80
20 kHz rate
New TI version 3
Linux CPU

## Slide 11: Parallel DAQ
Dead-time for N modules will be DN ~ (D1)N, D1 is dead-time for 1 module.
If D1 = 30% => D3 ~ 2%

## Slide 12: ADC test setup

## Slide 13: ADC test procedure
FPGA reprogrammed to generate programmable width ECL signals
Signal width ~ amplitude
12000 events generated the width changes every 1000 event
Each ADC channel should have 12 peaks with fixed amplitude 1000 events in each
Pedestal events also generated to look the width of the pedestal

## Slide 14: TDC test setup

## Slide 15: TDC test procedure
Sparsification CRL #1
Full scale time window = 400 ns
Sparsification threshold = 0 ns
Sparsification CRL #2
Full scale time window = 400 ns
Sparsification threshold = 300 ns
Sparsification CRL #3
Full scale time window = 496 ns
Sparsification threshold = 424 ns
0 ns
325 ns
400 ns
325 ns
300 ns
Common stop
Sparse. Thresh.
Sparse. Thresh.
0 ns
325 ns
400 ns
Common stop
Common stop

## Slide 16: Test results
200 TDC-s tested
175 good TDC-s
20 has problems in single channel
5 has other issues
120 ADC-s tested
97 good ADC-s
22 has problems in single channel
1 does not accept gate
All tested modules are marked and the DB created describing the location and test results
27 GAC-s tested
22 good GAC-s
28ATC-s tested
18 good ATC-s
19 SFI-s tested
18 good SFI-s

## Slide 17: Summary
Hardware
  TDC and ADC modules tested and ready to use
  6 power supplies available, 18 FASTBUS bins available
  12 FASTBUS crates are now installed in the platform in test-lab and powered with 4 power supplies.
  18 SFI modules available and tested
  22 GAC and 18 ATC cards available and tested
Software
  Sparsification mode tested with Intel CPU
  SFI sequencing tested have not seen an improvement in dead time
  Event blocking tested for 3 module 64 channels configuration, life time increased from 68% to 80% for 20 kHz generator rate.
  Parallel DAQ work in progress