## Slide 1: APEX DAQ rate capability
April 19th 2015
Alexandre Camsonne

## Slide 2: Outline
HRS
Channel count
Trigger rates
Triggers
Deadtime using 3 crates
TDC sparsicification
Sci Fi readout
Possible upgrades
  FADC
  TDC
  Event blocking

## Slide 3: HRS Detectors
Central Ray
2 PMTs
S2m
Gas Cherenkov
Total 4 layers
368 wires each
VDC
S0
16 Paddles
10 PMTs
2 segmented layers of lead glass blocks
RHRS
Calorimeter
S1 ?
Notes: Scintillators used for timing and triggers
Left HRS
34 blocks in each layer
15x15x30cm dimensions
Right HRS
48 blocks in preshower
10x10x35cm dimensions
75 blocks in shower
15x32.5x15cm dimensions
Used for PID
Right arm GC used as part of trigger to achieve pion rejection on trigger level

## Slide 4: Detector channels
Detector | Detector Channels | DAQ channels
Gas Cherenkov | 10 | 10 time 10 amplitude
Scintillators | 16+1 | 34 time 34 amplitude
Calorimeter | 128 – RHRS 68 – LHRS | 128 amplitude 68 amplitude
VDC | 1472 | 1472 time
RHRS
172 amplitude channels – 3 ADC modules
1516 time channels – 16 TDC modules
LHRS
112 amplitude channels – 2 ADC modules
1516 time channels – 16 TDC modules

## Slide 5: Trigger rate
Setting | A | B | C | D
Energy (GeV) | 2.2 | 4.4 | 1.1 | 3.3
e- (kHz) | 4500 | 700 | 6000 | 2900
π- (kHz) | 100 | 2200 | 36 | 1000
e+ (kHz) | 31 | 3.6 | 24 | 23
π+ (kHz) | 100 | 2200 | 36 | 1000
Trigger
Reduce coincidence timing
Suppress pions
Possible Improvements
DAQ
Using of sparsification

## Slide 6: Trigger Logic
Electron Arm Trigger (T1)
  Electron S2m
Positron Arm Trigger (T3)
  Positron S2m
Coincidence Trigger (T4)
  Electron S2m + Positron S2m
“Golden” Coincidence Trigger (T6)
  Electron S2m + Positron S2m + Positron Gas Cherenkov
Add muon trigger ?

## Slide 7: Trigger rate
Setting | A | B | C | D
Energy (GeV) | 2.2 | 4.4 | 1.1 | 3.3
e- (kHz) | 4500 | 700 | 6000 | 2900
π- (kHz) | 100 | 2200 | 36 | 1000
e+ (kHz) | 31 | 3.6 | 24 | 23
π+ (kHz) | 100 | 2200 | 36 | 1000
T6 rate (kHz) 20ns window π+ rejection = 30 | 3.2 | 4.5 | 3.0 | 4.4
For 4 kHz DAQ dead time is 10%
APEX can run without any improvement to DAQ!

## Slide 10: g2p

## Slide 11: 3 Crate configuration
RHRS
172 amplitude channels – 3 ADC modules
1516 time channels – 16 TDC modules
Crate 1: 5 TDC + 1 ADC
Crate 2: 5 TDC + 1 ADC
Crate 3: 6 TDC + 1 ADC
LHRS
112 amplitude channels – 2 ADC modules
1516 time channels – 16 TDC modules
    Crate 1: 5 TDC + 1 ADC
    Crate 2: 5 TDC + 1 ADC
    Crate 3: 6 TDC
By reducing number of channels which go to each crate we reduce the amount of data to be recorded and improve dead time.

## Slide 12: Sparsification
Common Stop
Start from background
Start from signal
full- scale window
0-32 µs (8ns step)
Sparsification
0-8µs
(0.5ns step)
Gate
Typical TDC spectrum without sparsification. Only 300ns window out of 1.5µs full scale window has useful data
Enabling of sparsification allows to significantly reduce event size. All TDC modules have been tested to work in sparsification mode.
To be implemented ( Kalyan )

## Slide 13: SciFi Readout
4 FADC in each HRS
1 VME64X crate in each HRS
  21 slots
  Intel CPU readout
  100 MB/s sustained rate ( 5 x faster than Fastbus )
Could add Cerenkov and other detectors if more FADCs

## Slide 14: 1881M
10 Megawords /s : 40 MB/s – 20 MB/s sustained
16 Hits/channel
Conversion time : 12 us 13 bit – 9 us 12 bit
AS-AK Handshake Time: 125 nsec typical, 150 nsec maximum. DS-DK Handshake Time: 65 nsec typical, 75 nsec maximum.

## Slide 15: 1877S
10 Megawords /s = 40 MB/s , 20 MB/s sustained
16 Hits/channel
Conversion time : 1.2 us + 50 ns/hit 1.75 us Min 7.6 us Max
Suppression : windows with steps of 8 ns
AS-AK Handshake Time: 125 nsec typical, 150 nsec maximum. DS-DK Handshake Time: 65 nsec typical, 75 nsec maximum

## Slide 16: ADC channels on FADC
RHRS
172 amplitude channels – 11 FADC modules
1516 time channels – 16 TDC modules
Crate 1: 5 TDC
Crate 2: 5 TDC
Crate 3: 6 TDC
Crate 4: 11+4 FADC fits in one crate
LHRS
112 amplitude channels – 7 ADC modules
1516 time channels – 16 TDC modules
    Crate 1: 5 TDC
    Crate 2: 5 TDC
    Crate 3: 6 TDC
    Crate 4: 7 + 4 FADC
Gain from 1.4 us up to 7.25 us encoding depending on VDC occupancy
Deadtime dominated by VDC data transfer
26 FADCs ~ 130 K$
Might be able to borrow from JLab

## Slide 17: Event size TDCs
4 x 368 = 1472 channels
368 = 4 TDCs
Cluster 5 wires = 2 + 5 = 7 words = 28 bytes
Time transfer crate block transfer :
  50 us + 2 /16 * 5 = 51 us
Time transfer standard :
  7 * ( 150 + 75 ) = 7 x 225 = 1.575 us
Block transfer saves time for more than 222 words

## Slide 18: TDCs on VME
3 VME64X crates available from DVCS
3 VXS crates
Need 1472 channels for each HRS
  CAEN 1190 : 12 TDCs = 120 K$
  VETROC : 12 VETROC = 60 K$
  F1 : 24 F1
Max 240 K$, can do 120 K$ with JLAB VETROC
Need 48 F1 ( most likely not enough )
Can mix F1 and VETROC
Improvement :
  x2 with 1 crate 9 KHz
  x4 with 2 crates 18 KHz

## Slide 19: Event Blocking
Readout
Triggers
Readout
Overhead
Readout
Overhead
Readout
Overhead
Overhead
Triggers
Readout
Readout
Readout
Readout
Overhead
For one channel readout.
4 events no EB – 280 µs
4 events with EB – 220 µs
20µs 50µs 20µs 50µs 20µs 50µs 20µs 50µs
20µs 50µs 50µs 50µs 50µs
Readout time is 50 µs + 2 µs per 16 channel.
64 channels – 58 µs
128 channels – 66 µs

## Slide 20: Event Blocking test results
Event Blocking | Number of channels | Number of modules | Life time (%)
1 | 1 | 1 | 71
4 | 1 | 1 | 90
1 | 64 | 3 | 65
4 | 64 | 3 | 80
20 kHz rate
New TI version 3
Linux CPU

## Slide 21: Event blocking
Can improve up to 30 % deadtime
  4.5 KHz to 5.8 KHz
Depends on rate and event size
Need 5 new TI for each HRS = 10 TI and new TS
If all VME and event blocking : up 200 KHz rate expected

## Slide 22: Conclusion
Using TDC sparsification should reach 4.5 KHz with 10 % deadtime
Use of 3 fastbus crates helps reduce deadtime
Will try implement event blocking but needs hardware , testing and modification of analysis software
Optimize transfer depending on event size
Options for upgrade for some costs