## Slide 1
Run 1087 (E=10 GeV)
electrons
pions
electron/pion separated by Cherenkov
>250 vs < 50
Notes: <number>

## Slide 2
Run 1087 (E=10 GeV): Cluster ADC

## Slide 3
Electron (E=1-10 GeV): Cluster ADC
E (GEV) | MEAN | SIGMA | S/M
1 | 2020 | 383 | 19%
2 | 3764 | 561 | 15%
6 | 11780 | 1288 | 11%
10 | 13990 | 1298 | 9.3%
16 | 6744 | 585 | 8.7%
1GeV
2GeV
6GeV
10GeV
16GeV

## Slide 4
Electron (E=1-10 GeV): Cluster ADC

## Slide 5
FTBF Position Measurements

## Slide 6
FTBF Position Measurements
Run Number | Energy (GeV) | Delta X (cm) | Delta Y (cm)
1064 | 6 | 0.192 | 0.564
1065 | 6 | 0.195 | 0.571
1066 | 10 | 0.313 | 0.544
1067 | 10 | 0.323 | 0.544
1082 | 10 | -0.220 | 0.587
1087 | 10 | -0.209 | 0.596
==================6 GeV====================
…
1064
1065
==================10 GeV====================
HV setting is:
HV for FADC channel 0 to 7:
SC1 SC2 Ecal1 Ecal2 Ecal3 PrSh1 PrSh2 PrSh3
-1260 -1280 -720 -820 -840 -1640 -1480 -1800
1066
1067
...
1074 Moved table from (1447, 306) to (1442,306), one spill only
…
1082 with events 278606
…
1087 with events 309246
Shift between (1064,1065) to (1066, 1067) could be due to gain change. Jixie will provide calibrated data
Shift between (1066,1067) and (1082-1087) consistent with motion table movement

## Slide 7
Run 1087 (E=10 GeV) - Electrons
MWPC
Preshower
ECal
all
ADC>10000

## Slide 8
Run 1087 (E=10 GeV) - Electrons
MWPC
Eff= (ADC>10000)/all
3.2mm

## Slide 9
Run 1087 (E=10 GeV): MWPC vs ECal
MWPC X
Ecal X
MWPC Y
Ecal Y

## Slide 10
Run 1087 (E=10 GeV): MWPC vs ECal
Delta X
Delta Y