Radiological Safety Analysis Document

This Radiological Safety Analysis Document (RSAD) will identify the radiation budget for the experiment, the verification process for the radiation budget, and controls with regard to production, movement, or import of radioactive materials.

I.	Description

Experiment E08-011, e −2H Parity Violating Deep Inelastic Scattering (PVDIS) at CEBAF 6 GeV, will be run in Hall A from approximately November 3, 2009 to December 23, 2009.  The current for this experiment is up to 100 microamps, with cryogenic D2, carbon, tantalum beryllium and aluminum targets, at energies of 4.870 and 6.000 GeV.  Experiment documentation can be found at http://hallaweb.jlab.org/experiment/E05-007/pac33/PR_pvdis3_submitted.pdf.

### II.	Summary and Conclusions

The experiment is calculated to use 116.8% of the annual design goal at the Jefferson Lab boundary for 1104 hours run-time.  There will be no need for manual manipulations of the target during the experiment. The experiment will be periodically monitored by the Radiation Control Department to ensure that the site boundary goal is not exceeded. The experiment will cause Radiation Areas and High Radiation Areas in the Hall.  Adherence to this RSAD is vital.

III.	Calculations of Radiation Dose at Site Boundary

The radiation budget for a given experiment is the amount of radiation that is expected at site boundary as a result of a given set of experimental conditions.  This budget may be specified in terms of mrem at site boundary or as a percentage of the Jefferson Lab design goal for dose to the public, which is 10 mrem per year.  The Jefferson Lab design goal is 10% of the DOE annual dose limit to the public, and cannot be exceeded without prior written consent from the Radiation Control Department Head, the Director of Jefferson Lab, and the Department of Energy.  This design goal is also referenced in the JLab Final Safety Assessment Document (FSAD).

The radiation budget for the PVDIS experiment (E08-011), with Spokespersons Xiaochao Zheng, Bob Michaels and Paul Reimer, is 11.68 mrem, which exceeds the Jefferson Lab annual design goal.  The attached spreadsheet details the calculations.  Historically, radiation budget estimates are known to be conservatively high, and based on this experience, it is likely that actual measured doses at the boundary will be considerably lower.  The experiment will proceed with specific measures for monitoring the boundary radiation dose as detailed below, and will be halted if it appears the annual design goal may be exceeded.  The Hall’s budget will be monitored during the experiment by using the active monitors at the Jefferson Lab site boundary to keep up with the dose for the individual setups.

In this case, we must consider both the existing budget expenditure thus far in the year, plus the anticipated dose from the current experiments.  Experiments in Hall A primarily impact the dose at boundary monitors RBM1 and RBM2.  As of October 5th, cumulative doses were 0.528 and 0.544 mrem at these locations, respectively.  The remaining budget for these locations is 9.472 and 9.456 mrem.  Also, RBM3 has a cumulative dose of 0.978 mrem as of the same date, but this monitor is primarily impacted by Hall C operations.

Cumulative doses at all three end station monitors will be assessed on a weekly basis upon starting PVDIS.  Cumulative dose “action levels” will be used as follows.

- Weekly updates will be provided (in writing, via email) to the Physics and Accelerator Division DSOs, indicating cumulative doses.

- If total annual dose on a monitor reaches 6 mrem, assessment frequency will be increased to twice weekly (Monday and Thursday), and a projection will be made based on accumulation rate, with a forecasted date for reaching 8 and 10 mrem.  If the forecast indicates a likelihood of exceeding 10 mrem at the existing rate, a meeting will be held with Physics Division Associate Director to assess options to reduce the dose rates or otherwise prevent exceeding the design goal.

- If measured dose reaches 8 mrem, a “hold” will be placed on the experiment, and a written plan will be produced specifying the intended path forward.

- If the intention is to exceed 10 mrem, written notification will be made to TJSO, otherwise, the experiment will be halted when the boundary dose reaches 9 mrem.

### IV.	Radiation Hazards

The following controls shall be used to prevent the unnecessary exposure of personnel and to comply with Federal, State, and local regulations, as well as with Jefferson Lab and the Experimenter's home institution policies.

A.  From Beam in the Hall

When the Hall status is Beam Permit, there are potentially lethal conditions present.  Therefore, prior to going to Beam Permit, several actions will occur.  Announcements will be made over the intercom system notifying personnel of a change in status from Restricted Access (free access to the Hall is allowed, with appropriate dosimetry and training) to Sweep Mode.  All magnetic locks on exit doors will be activated.  Persons trained to sweep the area will enter by keyed access (Controlled Access) and search in all areas of the Hall to check for personnel.

After the sweep, another announcement will be made, indicating a change to Power Permit, followed by Beam Permit.  The Run-Safe boxes will indicate "OPERATIONAL" and "UNSAFE".  IF YOU ARE IN THE HALL AT ANY TIME THAT THE RUN-SAFE BOXES INDICATE "UNSAFE", IMMEDIATELY PRESS THE “PUSH TO SAFE” BUTTON ON THE BOX.

Controlled Area Radiation Monitors (CARMs) are located in strategic areas around the Hall and the Counting House to ensure that unsafe conditions do not occur in occupiable areas.  This experiment may create regions of increased radiation outside the hall in areas not normally controlled for radiological purposes.  The RadCon Department will monitor the CARMs and make surveys as necessary to assess the impact of the experiment on radiation levels around the hall.

NOTE:

Calculated dose rates for this experiment are significantly higher than the design average dose rate goal.  Any indication that the levels may exceed 5 mrem/hr dose rate in an occupied area will require immediate mitigation, with continued operations contingent on a formal review of conditions and operational parameters, and final approval of operations exceeding this threshold by the Jefferson Lab RadCon Manager, in consult with Physics and Accelerator Division Safety Officers.

B.  From Activation of Target and Beamline Components and Other Materials in the Hall

1.  The Radiation Control Department shall be consulted for all movement of used targets, collimators, and shields.  The Radiation Control Department will assess the radiation exposure conditions and will implement controls as necessary based on the radiological hazards.

2.  There shall be no local manipulation of activated target configurations without direct supervision by the Radiation Control Department.  Remote movement of target configurations is permitted using appropriately reviewed and approved methods.

3.  No work is to be performed on beamline components, which could result in dispersal of radioactive material (e.g., drilling, cutting, welding, etc.).  Such activities must be conducted only with specific permission and control by the Radiation Control Department.

4.  The target chamber area and downstream beamline are expected to become significantly activated.  No work on this portion of the beamline is to be conducted without RCD review.

5.  Given the conditions for this experiment, it is expected that High Radation Areas will develop near the target area and beam dump/hall interface relatively quickly.  It is possible that these areas may contain dose rates > 1 rem/hr at 30cm.  If such conditions occur, special, positive access control procedures shall be used for all entries to the hall (even if the hall is in Restricted Access).  In addition, if general area dose rates in the hall exceed the threshold for Radiation Area posting (5 mrem/hr), the radiological status of the entire hall will be upgraded to that of “Radiation Area”, and work in the hall will be governed by specific RWPs.

NOTE: Work planning for all radiological work shall be coordinated through the hall work coordinator (E. Folts) using the ATLis work planning tool.

6.  This experiment is expected to produce low levels of airborne radioactivity which may impact environmental effluent standards and produce localized or generalized buildup of surface contamination in the hall.  Airborne radioactivity concentration in the hall is measured continuously.  If airborne radioactivity concentration as monitored by the AMS-4 air monitor in the experimental hall exceeds an average of 1.0E-6 microCurie/cm3 for a period of greater than 5 consecutive days, the Radiation Control Department (RCD) will require a meeting with the experimenters and the Physics Division Safety Officer to determine if the experimental conditions are accurate, and to assess what actions may reduce the airborne radioactivity levels to ensure that Jefferson Lab dose to the public from release of airborne radioactivity limits are not exceeded.

7.  Low levels of surface contamination are expected in some areas.  The RCD will monitor for the presence of this hazard as appropriate, and will require administrative controls and/or PPE commensurate with the conditions.  All posted guidance for contamination control must be observed.  Refer to the General Access RWP for details regarding controls for potentially affected systems.

C.  Other Sources

All radioactive materials brought to Jefferson Lab shall be identified to the Radiation Control Department.  These materials include, but are not limited to radioactive check sources (of any activity, exempt or nonexempt), previously used targets or radioactive beamline components, previously used shielding or collimators, or He-3 containers.  The RCD inventories and tracks all radioactive materials onsite.  The Radiation Control Department may survey the experimental setup before experiments begin as a baseline for future measurements if significant residual activity levels are present.

Tanks or cylinders of He-3 containing more than 10 mCi of tritium (H-3) shall not be stored or used in an experimental hall without the express, written permission of the RadCon manager.  Any containers of He-3 brought on site shall be assessed for the tritium content before use.  Additionally, He-3 containers should not be stored in the experimental hall when not in use.

V.	Incremental Shielding or Other Measures to be Taken to Reduce Radiation Hazards

No additional shielding is planned for this experiment.  Section III above outlines administrative controls intended to mitigate impact on the annual dose budget.

The RCD Head will notify the Hall Leader and Physics Division Safety Officer of any identified trends which might impact access to the hall or create conditions requiring broad changes to radiological working standards (i.e. General Access RWP revision).  The RCD head will recommend engineered or other controls considered necessary to prevent significant degradation of the radiological conditions in the hall.

VI.	Operations Procedures

1.  All experimenters must comply with experiment-specific administrative controls.  These controls begin with the measures outlined in the experiment's Conduct of Operations Document, and also include, but are not limited to, Radiation Work Permits, Temporary Operational Safety Procedures, and Operational Safety Procedures, or any verbal instructions from the Radiation Control Department.  A general access RWP governing access to the Halls and the accelerator enclosure must be read and followed by all participants in the experiment.  This RWP can be read and electronically signed online at: http://www.jlab.org/div_dept/train/Knowledge_Docs/GAPelec.pdf

2.  Any individual with a need to handle radioactive material at Jefferson Lab shall first complete Radiation Worker (RW-I) training.

3.  There shall be adequate communication between the experimenter(s) and the Accelerator Crew Chief and/or Program Deputy to ensure that all power restrictions on the target are well known.  Exceeding these power restrictions may lead to excessive and unnecessary contamination, activation, and personnel exposure.

4.  No target chamber or downstream component may be altered outside the scope of this RSAD without formal Radiation Control Department review.  Alteration of these components (including the exit beamline itself) may result in increased radiation production from the Hall and a resultant increase in site boundary dose.

5.  Any requested changes outside of the experimental parameters submitted for the calculation of the radiation budget (i.e., current, energy, target material, target thickness, run time) for this experiment shall require a formal review by the Radiation Control Department, and a new revision to the RSAD.

VII.	Decommissioning and Decontamination of Radioactive Components

Experimenters shall retain all targets and experimental equipment brought to Jefferson Lab for temporary use during the experiment.   After sufficient decay of the radioactive target configurations, they shall be delivered to the experimenter's home institution for final disposition.  All transportation shall be done in accordance with United States Department of Transportation Regulations (Title 49, Code of Federal Regulations) or International Civil Aviation Organization (ICAO) regulations.  In the event that the experimenter's home institution cannot accept the radioactive material due to licensing requirements, the experimenter shall arrange for appropriate funds transfers for disposal of the material.  Jefferson Lab cannot store indefinitely any radioactive targets or experimental equipment.

The Radiation Control Department may be reached at any time through the Accelerator Crew Chief (269-7045) or directly by calling the RadCon Cell Phone (876-1743). On Weekends, Swing Shift, and Owl Shift, requests for RadCon support should be made through the Crew Chief. This will ensure that there is prompt response with no duplication of effort.

Approvals:

______________________________		________________

Radiation Control Department Head			Date

RSAD E08-011

Attachment A