EC and SPD Updates
Xiaochao Zheng (University of Virginia) for
The SoLID EC Working Group
SoLID Collaboration Meeting
May 6-7, 2016

---
To Do List from Previous Collab Meeting
Determine the final design of module
Use paper sheets or not, what kind of paper / Lego Lock ?
Most materials are ready (awaiting WLS fiber), more modules can be assembled soon
Add real-time pressure and compression monitoring
How to let fibers go through module
Measure mechanical property of the scintillator
Cosmic test → Beam test
Resumed working with ANL/Chicago engineer on the Ecal support
Learn from other experiments
Funding application

---
More To do Items from 2015 Collab Meeting
Tests requiring beam → spring 2016
FASPD uniformity test
LASPD timing test
preshower prototype radiation resistance

---
More To do Items from 2015 Collab Meeting
Tests requiring beam → spring 2016
FASPD uniformity test
LASPD timing test
preshower prototype radiation resistance ― 8 preshower prototypes (4 from Kedi ― 科迪； 4 from CNCS ― 北京核仪器厂) radiated during the Hall A Spring 2016 run, need to measure their post-radiation performance ― summer 2016
both need tracking, maybe summer (if GEM is ready) or Fall 2016 (with beam)

---
http://hallaweb.jlab.org/experiment/E05-007/SoLID/EC/meetings/2014-test/2016-test/radHardness/SoLID_preshower_radhardness.html

---
On Ion Chamber
Upstream of Target Chamber

---
To Do List from Previous Collab Meeting
Determine the final design of module
Use paper sheets or not, what kind of paper / Lego Lock ?
Most materials are ready (awaiting WLS fiber), more modules can be assembled soon
Add real-time pressure and compression monitoring ― ongoing (SDU, THU)
How to let fibers go through module ― ongoing (SDU, THU)
Measure mechanical property of the scintillator
Cosmic test → Beam test
Resumed working with ANL/Chicago engineer on the Ecal support ― ongoing
Learn from other experiments
Funding application ― China, plan to apply for US NSF PIRE
Preliminary test of light yield
also studying pressure vs. temperature (-30,30C) (THU) and how to use spring washer to stablize the inner compression
― done (UVa, tensile, comparable to regular polysterene)

---
Progress – Some Pictures
SDU now have two modules (no fiber), one 170-layer with inner sensor for mechanic testing; one full module
Kedi TiO2-based coating (ALICE BC622A; LHCb: IHEP chemically-treated )
THU also have two modules (no fiber), one for compression testing; one full module
UVa: no module,small tests; sending Y11 fibers to China; giving weekly instructions; also FMPMT paper accepted by NIM

---
Module Load Calculation ― from Sept. 2015
Turn nuts further so weight is completely balanced by static friction (642 lbf load on 6 rods, factor 2 included)
500kg force to flatten layers; put in rods, turn nuts to snug; 78 lbf load on 6 rods after assembly plates are removed
Cantilevering add more stress to top 3 rods, 400 lbf alone on top rod (2 included)
cantilevering on back support
Do not need support between preshower and shower ???
Preloading on assembly stand
       loading
layers
layers
layers

---
Module Load Calculation ― Progress
Turn nuts further so weight is completely balanced by static friction (642 lbf load on 6 rods, factor 2 included)
500kg force to flatten layers; put in rods, turn nuts to snug; 78 lbf load on 6 rods after assembly plates are removed
Cantilevering add more stress to top 3 rods, 400 lbf alone on top rod (2 included)
cantilevering on back support
Preloading on assembly stand
       loading
500kg achieved at both SDU and THU;
procedure works mostly, but inner compression is half of the outer compression value after removing compression plates. (SDU result, brass rods, using inner sensors) (Note this partial reduction is expected and determined by the Young's modulus of all material)
calc based on 0.1 static friction. Actual value is 0.1 if using Tyvek, or 0.2 if using printer paper (measured at SDU), affects light yield
need to establish an relation between torque wrench and inner compression, since inner sensors won't be available (ongoing)
need mechanical test to measure minimum compression for friction-support (SDU/THU ongoing)
Two planes, one for preshower, one for cantilevering Shower.
“If back-support plane alone works for the Shower, it would be the simplest solution. So we start from here, until there is a show stopper.”
need stainless steel rods and at least the top rod needs to be 3mm dia (current 2.5mm)
Need full mechanical tests to finalize. Decision on the use of side plates (ALICE) and a 3rd support plane between PS and Shower both depend on these test outcome.
layers
layers
layers

---
assume 10cm fiber bending radius (no loss)
ECal Support Design (4/21/2016)
nuts inside hold compression
rigid spacer

---
ECal Support Design (4/21/2016)
need improvement/testing to minimize module sagging (currently 2.9mm if cantilevered by rods only ― SDU results)

---
Light Yield Test ― “hedgehog” test
LHCb used “hedgehog” test for screening the scintillator plate quality, but there is no detailed picture/diagram for the setup. We built our own “hedgehog” setup at UVa:
alternating layers of scintillator and reflective material on a plastic holder
96 Y11(200)MS fibers “sticking out” from holes (no mirrored end)
fibers glued to a permanent cylinder (with holes) and coupled to XP2262 PMT through optical grease
loose reflective layer on the tile edge.

---
single p.e.
single p.e.
main
Caveat: All tests done without lead sheets, so “no reflector” is not the same as “no reflector in shashlyk module”
printer paper with loose edge wrapping:
34p.e./25 layers = 1.36p.e./layer
no reflector:
39p.e./25 layers = 1.56p.e./layer
Tyvek with loose edge wrapping:
36p.e./25 layers = 1.44p.e./layer
(Also Tyvek w/o edge wrapping: 28.5p.e/25 layers = 1.14 p.e./layer, no picture yet)

---
Light Yield Test ― side readout (SDU)
The reflectors are changed everyday: print paper, aluminum, Tyvek, or nothing.
Aluminum plate
lead
reflector
scintillator
reflector
PMT
Currently, about 10 p.e./layer, could be consistent with the hedgehog test considering the 10% fiber trapping efficiency (multi-cladding).
Tyvek slightly higher than printer paper and aluminum (these two are comparable)

---
Light Yield ― Projection
this is 0.36p.e./MeV, compare to:
LHCb: 2.6-3.5 p.e./MeV (4mm sci, 2mm lead)
ALICE: 4-4.4 p.e./MeV (1.76mm sci, 1.44mm lead)
KOPIO: 53 p.e./MeV (0.275mm lead, 1.5mm sci, no paper, no optical contact between sci and lead; APD)
(all similar fiber density, Y11(200)MS)

---
Light Yield ― Projection
Additional #2: Our Tyvek is from LOWES (with blue prints). Anyone knows where to find pure white Tyvek 1055B?
Additional #1: LHCb simulation reported *3 light yield for mirrored tile edge vs. transparent edge, and *2 for chemically-treated edges (diffusive reflection, can also use Bicron TiO2 paint BC622A) (ordering BC622A now)
Assuming: for 1 GeV electron, 0.2 GeV is Edep in scintillators (actually sampling may be different ― need beam test)
From Preshower test, but LHCb reported similar light yield (ordering BCF91A-MC now, but won't ship until Sept.)
specific to SoLID
If taking the optimistic direction in all cases, our light yield should be within 20% or factor two of other experiments (quality of Tyvek?), and may reach 3.2%

---
Light Yield ― Projection
If taking the pessimistic direction in all cases, we will
* use Y11 instead of BCF91A ($200k → $600k)
* eliminate clear fiber (-$500k) and try APD or SiPM readout ― APD has S/N ratio (low gain, ~100); SiPM has radiation issue. ― cost to be studied

---
Plan
SDU and THU will continue with mechanical testing, providing inputs to support design
SDU and THU will assembling full modules and carry out cosmic (horizontal) test. Vertical test will be ideal but hard to do. (within 2016?)
UVa hedgehog test:
white Tyvek; BCF91A; TiO2 paint on edge; add lead layers; (hopefully by early fall 2016)
PMT base tests? – need to follow up with Jlab detector group;
After all these tests, hopefully we will have a better idea on:
choice of reflective material
light yield
module infrastructure (compression, etc.)
Then we can talk about beam tests for the shashlyk modules;
Fall 2016 parasitic or cosmic+GEM tests of FA and LASPDs
Awaiting PIRE solicitation

---
Pre R&D Need – end of 2015
At least 1/2 postdoc to “develop an end‐to-end realistic simulation and reconstruction to further optimize cost and physics reach and derive clear performance requirements for the individual subdetectors.”
Resources to build 4 more prototype modules, PMTs and HVs, combine with the SDU prototypes and conduct in-beam test:
construction: $34k material + ($10k-$20k) assembly stand, 1/4 postdoc, 1/2 tech or grad students, 3 summer undergrad.
testing: partial postdoc, 1/2 grad student
Other items underway and covered by UVa
FASPD uniformity test
LASPD timing test with GEM
radiation resistance test
continue working with ANL engineers on module&support
continue working with JLab detector group on PMT base design&testing
misc: fiber, fiber connector, … ...

---
Pre R&D Need – end of 2015
At least 1/2 postdoc to “develop an end‐to-end realistic simulation and reconstruction to further optimize cost and physics reach and derive clear performance requirements for the individual subdetectors.”
Resources to build 4 more prototype modules, PMTs and HVs, combine with the SDU prototypes and conduct in-beam test:
construction: $34k material + ($10k-$20k) assembly stand, 1/4 postdoc, 1/2 tech or grad students, 3 summer undergrad.
testing: partial postdoc, 1/2 grad student
Other items underway and covered by UVa
FASPD uniformity test
LASPD timing test with GEM
radiation resistance test
continue working with ANL engineers on module&support
continue working with JLab detector group on PMT base design&testing
misc: fiber, fiber connector, … ...
THU has picked up these 4 remaining modules. Still need to cover WLS fiber+PMT cost and material for the hedgehog test.
If light yield remains an issue, will need to order IHEP prototypes

---
Pre R&D Need – present
At least 1/2 postdoc to “develop an end‐to-end realistic simulation and reconstruction to further optimize cost and physics reach and derive clear performance requirements for the individual subdetectors.”
Resources to build 4 more prototype modules, PMTs and HVs, combine with the SDU prototypes and conduct in-beam test:
construction: $34k material + ($10k-$20k) assembly stand, 1/4 postdoc, 1/2 tech or grad students, 3 summer undergrad.
testing: partial postdoc, 1/2 grad student
Other items underway and covered by UVa
FASPD uniformity test
LASPD timing test with GEM
radiation resistance test
continue working with ANL engineers on module&support
continue working with JLab detector group on PMT base design&testing
misc: fiber, fiber connector, … ...
We have always assumed IHEP will make the Shashlyk modules for us. But communication has been difficult and political issues persist (also price inflation everytime I contact IHEP). Now it looks promising for the Chinese SDU+THU groups to make the 8 prototype modules (and may save some $$$ at mass production), but need significant pre-R&D and R&D and fund to buy material. The ECal design is still depending on these test results. ECal pre R&D should be one of the top priorities.