## Slide 1: THU2 cosmic test results (6)
Chendi Shen
2017.9.21

## Slide 2
Horizontal Test

## Slide 3
Npe=[(QDC-ped)*0.1*10-12]/(1.6*10-19*3.1*106)
Npe=92.4
HV=1000V scintillator for trigger(up: 10cm*5cm*5cm down: 3cm*3cm*3cm)
Total Number of events 805
No attenuator
Npe
Experiment 01
ECal
ECal
ECal
Move forward by 10cm
Move backwards by 10cm

## Slide 4
Npe=[(QDC-ped)*0.1*10-12]/(1.6*10-19*3.1*106)
Npe=90.1
HV=1000V scintillator for trigger(up: 10cm*5cm*5cm down: 3cm*3cm*3cm)
Total Number of events 487
No attenuator
Npe
Experiment 02
ECal
ECal
ECal
Move forward by 10cm
Move backwards by 10cm

## Slide 5
Npe=[(QDC-ped)*0.1*10-12]/(1.6*10-19*3.1*106)
Npe=103
HV=1000V scintillator for trigger(up: 10cm*5cm*5cm down: 3cm*3cm*3cm)
Total Number of events 428
No attenuator
Npe
Experiment 03
ECal
ECal
ECal
Move forward by 10cm
Move backwards by 10cm

## Slide 6
ECal
Move forward by 10cm
ECal
Move backwards by 10cm
Npe=90.1
ECal
Npe=92.4
Npe=103
10cm
10cm
The farther away from PMT, the smaller the NPE
The efficiency of the detector in horizontal test all is about 93%

## Slide 7
Gain of the PMT (SPE)
Because we want to complete the scale of the gain and continue the cosmic ray test as soon as possible, so we finished this experiment at the IHEP in Beijing.
Their research is development of the PMT for Jiangmen neutrino experiment, so they have a wealth of experience with the PMT gain scale

## Slide 8
PMT
LED
Signal generator
Amplifier (x11)
FADC (DT5751)
PC
Trigger (300HZ)
Dark Box

## Slide 9
Calibration of Amplifier
After adding the amplifier (110mv)
Before adding the amplifier（10mv）
x11
Signal generator
Oscilloscope
Signal generator
Amplifier
Oscilloscope

## Slide 10
Gain=[(p6-p1)*20*10-15]/(1.6*10-19)
=2.2*106
Gain(sp/sk)=3.1*106
Events: 66683
Events: 13820
Eff13820/80503=17%
P( 0, 0.17) = 84%
P( 1, 0.17) = 14%
1-P( 0, 0.17)-P( 1, 0.17) = 2%
P(SPE)/P(PE) = 14%/(14%+2%)
= 87.5%
1000V
Gaussian 1
Gaussian 2
Exponential
Gain

## Slide 11
Gain=[(p6-p1)*20*10-15]/(1.6*10-19)
=4.5*106
Gain(sp/sk)=6.5*106
Events: 64979
Events: 9257
Eff13820/80503=13%
P( 0, 0.13) = 88%
P( 1, 0.13) = 11%
1-P( 0, 0.13)-P( 1, 0.13) = 1%
P(SPE)/P(PE) = 11%/(11%+1%)
= 91.6%
1100V
Gaussian 1
Gaussian 2
Exponential
Gain

## Slide 12
Summary
The farther away from PMT, the smaller the NPE
The efficiency of the detector in horizontal test all is about 93%
The gain is obtained by NPE method. 2.2*106 (1000V), 4.5*106 (1100V)
Next to do
Now we have back to the vertical test and will get the Npe at 900V and 1000V.