## Slide 1
THU2 Cosmic test with SDU PMT
Chendi Shen
2017.11.16

## Slide 2
SDU PMT and SDU base

## Slide 3
1050V (total number of events: 550) npe=680

## Slide 4
1100V (total number of events: 50)
Npe=622
The speed of data taking is very slow (2 events/hour)
I will use one week to get 500 events

## Slide 5
 | THU PMT |  | SDU PMT
HV | hamamatsu | IHEP | SDU
900 | 1.3*106 |  | 
1000 | 3.1*106 | 2.2*106 | 
1050 |  |  | 4.11*106
1100 | 6.5*106 | 4.5*106 | 5.997*106
1200 | 13.3*106 |  | 11.227*106
1300 | 26.1*106 |  | 
Gain of PMT

## Slide 6
Test results
Chart: Number of photo electron with different Gain
| 900 | 1000 | 1050 | 1100 | 1200 | 1300 | 1400
hamamatsu（THU PMT） | 750 | 748 | 432 | 271 | 164 | 96
IHEP (THU PMT) | 1054 | 624
SDU (SDU PMT) | 680 | 622
The SDU PMT is saturated at 1100V, we don’t know whether it is also saturated at 1050V
We need to get Gain of the SDU PMT under 1000V and 900V to compare with the results what we have got.