## Slide 1: Shashlik module and PMT test
Ye Tian
02/01/2018
Shandong Univesity

## Slide 2: Outline
PMT in THU#2 linearity test
THU powder paint test
SDU#3 test with higher voltage
Shashlik cosmic ray test in JLab

## Slide 3: SDU PMT(67748) in THU#2 linearity test(20ns LED pulse) HV: 1100V(5.997*10^6)
300
Photon number
Charge(pC)

## Slide 4: HV: 1000V(2.9*10^6)
550
Photon number
Charge(pC)

## Slide 5: HV: 900V(1.295*10^6)
1100
Photon number
Charge(pC)

## Slide 6: The max current(calculated from amplitude)
900V
1000V
1100V
20mA
24mA
26mA

## Slide 7
Problems in linearity test:
All HVs has the max charge 250pC(-5%), the max NPEs are only influenced by the gain
For 1100V, 300 photon electrons would exceed linearity range, not corresponding to Chendi’s result
The width of LED pulse will influence the linearity range
The problems will be check soon

## Slide 8: THU powder paint reflectivity test
Receive three layers of powder paint from THU, test with similar test setup as previous five layers test(only 4 scintillator layers in test now).
Test setup
comparing with same setup that only powder paint layer replaced with Tyvek-lead-tyvek
Lead
Tyvek
Scintillator
Power paint layer

## Slide 9: Test Result
Material | PMT | HV(V) | Gain*10^6 | Charge | NPE
powder paint | 67702 | 1100 | 5.40 | 40.1 | 46.3
tyvek | 67749 | 1000 | 5.12 | 27.8 | 33.9
Assuming:
Power paint reflectivity =b
Tyvek reflectivity =a
We get:
So: b/a = 1.475
Which improve much comparing with SDU modules.
powder paint
tyvek

## Slide 10: SDU#3 Higher HV test
HV | Gain(*10^6) | Charge(pC) | NPE
900 | 1.35 | 106.1 | 491.2
950 | 2.01 | 155.4 | 483.2
1000 | 2.93 | 223.9 | 477.6
1100 | 5.4 | 427.3 | 494
1200 | 9.995 |  | 
1100V
1200V
Data may have problem

## Slide 11: Shashlik signal in JLab cosmic test
Shashlik Module | HV(start from run 1140) | Gain(10^6）
THU | 1700 | ?
SDU#1 | 1453 | 5
SDU#2 | 1294 | 5
THU(low signal)
SDU#1
SDU#2
fadc channel
fadc channel
fadc channel