## Slide 1
The attenuation length of the irradiated fiber
Dong Liu
Ye Tian
Mengjiao Li
2021.04.29

## Slide 2
The relation between the light intensity I and the distance the light travels in the medium x is ）, λ is the attenuation length of the medium
Comparing the radiated fiber with the non-radiated one :
+ )=ratio
Where, x is the fiber length, λ1 is the attenuation length of the irradiated fiber, λ2 is the attenuation length of the clear fiber without irradiation.
  PMMA : λ2=17.24
  BCF98-SC : λ2=4.55
Actually, , which will be described later

## Slide 3
 | Total radiation dose（MeV/cm^2） | PMMA (2m) | BCF98-SC (3m) | BCF98-SC (6m)
Sample 1 | 8.569E+11 | 0.990 | 0.979 | 0.978
Sample 2 | 1.360E+12 | 0.978 | 0.963 | 0.933
Sample 3 | 2.807E+12 | 0.967 | 0.938 | 0.919
Sample 4 | 3.665E+13 | 0.946 | 0.906 | 0.782
Comparison of relative light yield of clear fiber
 | Total radiation dose（MeV/cm^2） | PMMA (2m) 17.24m | BCF98-SC (3m) 4.55m | BCF98-SC (6m) 4.55m
Sample 1 | 8.569E+11 | 15.87 | 4.41 | 4.47
Sample 2 | 1.360E+12 | 14.47 | 4.30 | 4.32
Sample 3 | 2.807E+12 | 13.37 | 4.15 | 4.28
Sample 4 | 3.665E+13 | 11.66 | 3.96 | 3.83
The attenuation length of the irradiated fiber
Should be a constant for each type

## Slide 4
Light yield calculation:
is the fiber length, d is the radiation dose
We don’t know the function of d infludence, so written as f(d). Need accurate measurement and fitting to acquire f(d).
Intrinsic attenuation
Attenuation from
radidation

## Slide 5
 |  | f(d) |  | 
 | Total radiation dose（MeV/cm^2） | PMMA (2m) 17.24m | BCF98-SC (3m) 4.55m | BCF98-SC (6m) 4.55m
Sample 1 | 8.569E+11 | 0.00501 | 0.00698 | 0.0039
Sample 2 | 1.360E+12 | 0.0111 | 0.0128 | 0.0117
Sample 3 | 2.807E+12 | 0.1679 | 0.0212 | 0.013
Sample 4 | 3.665E+13 | 0.0278 | 0.0327 | 0.0413
The function f(d) of the irradiated fiber