Jefferson Lab Alignment Group

DATA TRANSMITTAL

TO:	Arun Saha, John LeRose					DATE: 12 Apr 1999

FROM:    Chris Curtis	                      			Checked:		# :  515

DETAILS:

Below are the results of the survey performed on the Hall A superharp /BPM assemblies on 5th April. This survey was carried out concurrently with the survey of the electron spectrometer. The numbers shown represent the amount (in mm) the component would have to 	move to be at the ideal location. A negative X is a movement to the beam right, and a positive Y is up. The results of the previous survey (3/31/99, Data Transmittal #511) are given for reference.

April 5				    X		   Y

SH 1				  0.16		-0.01

SH 2				  0.03		 0.00

March 31			    X		   Y

SH 1				  0.44		 0.04

SH 2				  0.19		-0.09