Jefferson Lab Alignment Group

DATA TRANSMITTAL

TO:	Javier Gomez, John LeRose, Arun Saha			DATE: 6 July, 1999

FROM:    Kelly J. Tremblay					Checked:	       	 # 536

DETAILS:

Below are the results from the spectrometer pointing surveys carried out May 28th, June 1st and June 2nd, 1999. The first and second surveys were hadron spectrometer locations.  The June 2nd survey was of the electron spectrometer concurrently with a superharp survey.  The angle results are.

Hadron - H052899A 	12.5

Hadron - H060199A 	12.5

Electron – E060299A	-12.5

================ RESULTS ======= H052899A - Hadron 1999.May.28

The central ray of the spectrometer is at   12.564 degrees.

It is missing the defined target center by   3.28 mm downstream,

and    -.42 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at   12.542 degrees.

To achieve this optimal setting,

make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent    3.34 mm downstream.

Vertical corrections (positive = up):

Central jack:            1.65 mm.

Upstream rear jack:      2.49 mm.

Downstream rear jack:    3.31 mm.

*.3dd = 0.91

*.9pr  = 0.00015

================ RESULTS =======  H060199A - Hadron 1999.June.01

Hadron check of the work from May 28th survey (different software)

The central ray of the spectrometer is at   12.563 degrees.

It is missing the defined target center by   3.23 mm downstream,

and    -.71 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at   12.541 degrees.

To achieve this optimal setting,

make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent    3.29 mm downstream.

Vertical corrections (positive = up):

Central jack:            1.85 mm.

Upstream rear jack:      2.45 mm.

Downstream rear jack:    3.57 mm.

*.3dd = 1.35

*.9pr = 0.00016

================ RESULTS ======= E060299A – Electron 1999.June.02

The central ray of the spectrometer is at  -12.526 degrees.

It is missing the defined target center by   1.26 mm   upstream,

and    -.67 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at  -12.534 degrees.

To achieve this optimal setting,

make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent    1.28 mm   upstream.

Vertical corrections (positive = up):

Central jack:            1.76 mm.

Upstream rear jack:      2.52 mm.

Downstream rear jack:    3.22 mm.

*.3dd = 1.41

*.9pr = 0.00013

Super Harp Movements in the Beam Following Coordinate System (E060299A)

Target      dz(mm)      	dx(mm)      	dy(mm)

SH1A        0.04      		-0.25      	 0.05

SH1B        0.26      		-0.18      	 0.10

SH1C        0.45      		-0.18      	 0.21

SH2A        0.20      		-0.10       	 0.06

SH2B        0.28       		-0.07      	 0.05

SH2C        0.37      		-0.03       	 0.11

The numbers shown represent the amount (in mm) the component would have to move to be at the ideal location (positive Z is downstream, positive X is beam left, positive Y is up). Tooling ball 'a' is beam right, 'b' is center, and 'c' is beam left.