Jefferson Lab Alignment Group

DATA TRANSMITTAL

TO:	Ed Folts, Javier Gomez, John LeRose			DATE: 30 Nov, 1999

FROM:    Chris Curtis						Checked:	       	 # 565

DETAILS:

Below are the results from the spectrometer pointing surveys carried out on the  		Hadron arm on November 19th and 30th.

================ RESULTS =============== H111999B

The central ray of the spectrometer is at   30.139 degrees.

It is missing the defined target center by    .14 mm   upstream,

and    -.34 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at   30.140 degrees.

To achieve this optimal setting, make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent     .14 mm   upstream.

Vertical corrections (positive = up):

Central jack: 1.14 mm.    U/s rear jack: 1.61 mm.    D/s rear jack:    2.30 mm.

*.3dd = 1.39     *.9pr = 0.13

================ RESULTS ===============  H113099A

The central ray of the spectrometer is at   45.200 degrees.

It is missing the defined target center by    .19 mm   upstream,

and     .25 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at   45.201 degrees.

To achieve this optimal setting, make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent     .20 mm   upstream.

Vertical corrections (positive = up):

Central jack:   .48 mm.      U/s rear jack:  .51 mm.      D/s rear jack:    1.94 mm.

*.3dd = 1.35	   *.9pr = 0.10