Jefferson Lab Alignment Group

DATA TRANSMITTAL

TO:	Ed Folts, Javier Gomez, Arun Saha, John LeRose	DATE: 3 Dec, 1999

FROM:    Chris Curtis						Checked:	       	 # 567

DETAILS:

Below are the results from the spectrometer pointing surveys carried out on the  		Hadron and Electron arms on December 2nd and 3rd.

================ RESULTS ===============  H120299A

The central ray of the spectrometer is at   16.553 degrees.

It is missing the defined target center by   2.92 mm downstream,

and    -.40 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at   16.533 degrees.

To achieve this optimal setting, make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent    2.97 mm downstream.

Vertical corrections (positive = up):

Central jack:   1.58 mm.       U/s rear jack:   2.39 mm.      D/s rear jack:   3.17 mm.

*.3dd = 1.41		*.9pr = 0.13

================ RESULTS ===============  E120399A

The central ray of the spectrometer is at  -16.522 degrees.

It is missing the defined target center by   2.56 mm   upstream,

and     .14 mm vertically (positive = up).

If the offset is corrected by secondary alignment, the

spectrometer will be at  -16.540 degrees.

To achieve this optimal setting, make the following adjustments:

Horizontal corrections:

Move rear jacks along tangent    2.60 mm   upstream.

Vertical corrections (positive = up):

Central jack:   1.39 mm.   U/s rear jack:   2.78 mm.   D/s rear jack:   3.11 mm.

*.3dd = 1.10		*.9pr = 0.11