| He METER | PROBE LENGTH | HI/TRIP POINT | ANALOG OUT | INTERVAL/LO/B/A |  |
| --- | --- | --- | --- | --- | --- |
| LQ1 He | 6” | 20% | 0-10V | N/A |  |
| LQ2 He | 20” | 60% | 0-10V | N/A |  |
| LQ3 He | 20” | 60% | 0-10V | N/A |  |
| LD1 He | 10” | 50% | 0-10V | N/A |  |
| RQ1 He | 6” | 20% | 0-10V | N/A |  |
| RQ2 He | 20” | 60% | 0-10V | N/A |  |
| RQ3 He | 20” | 60% | 0-10V | N/A |  |
| RD1 He | 10” | 50% | 0-10V | N/A |  |
| N2 METER | OVERALL LENGTH | ACTIVE LENGTH | ANALOG OUT | DIAMETER |  |
| LQ1 N2 | 38” | 18.35” | 4-20MA | 3/8” |  |
| LQ2 N2 | 24.25” | 11.75” | 0-10V | 3/8” |  |
| LQ3 N2 | 24.25” | 10 5/8” | 0-10V | 3/8” |  |
| LD1 N2 | 29.25 | 9 | 0-10V | 3/8” |  |
| RQ1 N2 | 38” | 18.35” | 4-20MA | 3/8” |  |
| RQ2 N2 | 24.25” | 11.75” | 0-10V | 3/8” |  |
| RQ3 N2 | 24.25” | 10 5/8” | 0-10V | 3/8” |  |
| RD1 N2 | 29.25 | 9 | 0-10V | 3/8” |  |
| DIPOLE He LVL LOWER PROBES | PROBE 1 | PROBE 2 | PROBE 3 | PROBE 4 | PROBE 5 |
| LD1 | 9 | 23.8 | 23.8 | 23.8 | 23.8 |
| RD1 | 23.9 | 24 | 24 | 24 | 24 |

The He level meters do not use the interval setting, nor do they use the set points “LO” “A” or “B”.  The trip point is set with the HI level relay contacts (HI set point on the front of the meter).  The length must be set from the front panel, this is only way to calibrate the probe to the meter. The only set up required for the N2 meters is the analog out put, probe/meter calibration is accomplished with a “wet cal” procedure.  The relay outputs are not used on the N2 meters and the Dipole lower He level probes.  The Dipole lower He level probes are a local readout only.