fa
	Operational Safety Procedure Form / (See ES&H Manual Chapter 3310 Appendix T1 Operational Safety Procedure (OSP) and Temporary OSP Procedure for instructions.)

Title:	BigBite GRINCH for the SBS experiments
Location:		Hall A, on the BigBite detector stack					Type:	● OSP / ◌ TOSP
Risk Classification / (per Task Hazard Analysis attached) / (See ES&H Manual Chapter 3210 Appendix T3 Risk Code Assignment.)				Highest Risk Code Before Mitigation				3
				Highest Risk Code after / Mitigation (N, 1, or 2):				1
Owning Organization:			Hall A		Date:	08/11/21
Document Owner(s):			Jack Segal, Bradley Yale, Eric Fuchey

DEFINE THE SCOPE OF WORK
Purpose of the Procedure – Describe in detail the reason for the procedure (what is being done and why).
This document describes the GRINCH detector which will be used in the BigBite spectrometer for the SBS experiments. This detector will be sitting in the BigBite detector stack in Hall A and is produced by the College of William and Mary.
Scope – include all operations, people, and/or areas that the procedure will affect.
Operation of the GRINCH in the Hall A SBS experiments. The scope of this OSP encompasses electrical issues associated with the GRINCH, work conducted under Covid-19 elevated MEDCON levels, access to the detector by use of a ladder, and opening the detector for maintenance or repair..
Description of the Facility – include building, floor plans and layout of the experiment or operation.
The GRINCH will be installed in the Bigbite detector stack, sitting between the two stacks of GEM planes (see illustration on the left). This detector stack is located in Hall A, on the BigBite platform. GRINCH is composed of a tank filled with C4F8 gas for Cherenkov / radiator. The light radiated by the charged particles is reflected by mirrors and collected by a matrix of 510 photomultiplier tubes. The signal from the PMTs are processed by NINO front-end cards, and are conveyed to the GRINCH DAQ located in the DAQ bunker in Hall A through 17-pair flat cables. In order to mitigate the signal attenuation along the 100 ft cables, an LVDS to ECL converter is also required in the hall between the detector package and the DAQ bunker. At the DAQ bunker there is a corresponding ECL to LVDS converter to convert the signals back to the correct format. The high voltage power supplies for the PMTs are also sitting in the DAQ bunker, and also requires cables running from the DAQ bunker to the detector stack. The low voltage for the NINO cards is located in the DAQ electronics bunker connected via 4 AWG welding cables to the low voltage distribution on the BigBite frame.

ANALYZE THE HAZARDS and IMPLEMENT CONTROLS
Hazards identified on written Task Hazard Analysis
Electrical shock, pressurized gas containers, oxygen displacement, potential for ladder work and work under Covid-19 elevated MEDCON.
Authority and Responsibility:
	Who has authority to implement/terminate
	Hall A/C leader, Hall A work coordinator, Todd Averett, Bradley Yale, Jack Segal
	Who is responsible for key tasks
	Todd Averett, Bradley Yale
	Who analyzes the special or unusual hazards including elevated work, chemicals, gases, fire or sparks (See ES&H Manual Chapter 3210 Appendix T1 Work Planning, Control, and Authorization Procedure)
	Jessie Butler (Hall A Work Coordinator), Jack Segal (Halls A and C Spectrometer Support Manager)
Personal and Environmental Hazard Controls Including:
	Shielding
	N/A
	Barriers (magnetic, hearing, elevated or crane work, etc.)
	GRINCH may need to be accessed by ladder. Safety training will be adhered and working in pairs will be advisable.
	Interlocks
	N/A
	Monitoring systems
	The high voltage will be monitored with the high voltage GUI. The gas will be monitored using a pressure gauge. Low voltage will be monitored with the low voltage GUI.
	Ventilation
	Standard Hall A ventilation. Local forced ventilation inside the detector if it is opened after filling. A buddy rule is instituted when opening the detector after it has been filled, minimum of two people required.
	Other (Electrical, ODH, Trip, Ladder) (Attach related Temporary Work Permits or Safety Reviews as appropriate.)
	Use of current limited high voltage supply at 2kV. Use of shielded HV cables and connectors. Use of current limited low voltage supply (12.5V, 120A). Gas supplied through a pressure regulator attached to the gas bottle with flow limited by a flow meter, restriction, and pressure relief valve.
List of Safety Equipment:
	List of Safety Equipment:
	N/A
	Special Tools:
	N/A
Associated Administrative Controls
	Setup, removal, or changes to the BigBite timing hodoscope setup may be coordinated through Todd Averett, Bradley Yale
Training
	What are the Training Requirements (See List of Training Skills)
	Hall A walk through. Radiation Worker I. ODH training. Ladder safety training, if requiring access by ladder. Pressure system training. ESC001, ESC003, ESC004, ESC007, and ESC008, current electrical training, subject to change as the new training is taking hold at the lab. Equipment specific training

DEVELOP THE PROCEDURE
Operating Guidelines
Normal operation of the system high voltage is documented in the HV “how-to”. Any other non-routine operation on the system shall not be made unless authorized by an individual in Section 5 and with training as noted above.
Notification of Affected Personnel (who, how, and when include building manager, safety warden, and area coordinator)
Contact Hall Work Coordinate prior to start of work
List the Steps Required to Execute the Procedure: from start to finish.
GRINCH installation prior to start of experiments in the BigBite detector stack. / Install GRINCH HV and data acquisition in the SBS DAQ bunker / Connect the electronics, data acquisition and high voltage. / Connect the gas system to the GRINCH / Gain match the PMTs, which requires access to the front end NINO cards (ladder / lift required)
Back Out Procedure(s) i.e. steps necessary to restore the equipment/area to a safe level.
Turn off high voltage and low voltage / Stop gas flow / Reassess the job before turning power back on
Special environmental control requirements:
	List materials, chemicals, gasses that could impact the environment (ensure these are considered when choosing Subject Mater Experts) and explore EMP-04 Project/Activity/Experiment Environmental Review below
	N/A
	Environmental impacts (See EMP-04 Project/Activity/Experiment Environmental Review)
	N/A
	Abatement steps (secondary containment or special packaging requirements)
	N/A
Unusual/Emergency Procedures (e.g., loss of power, spills, injury, fire, etc.)
In the event of injury, or an immediate emergency exists, call 911 and also notify: /  Guards (x5822) /  Occupational Medicine (x7539) /  Crew Chief (x7045) (if inside the fence) / In case of an injury follow standard JLAB procedures. Initial response cards are located with each phone for appropriate emergency phone numbers. Additional information can be found at https://jlabdoc.jlab.org/docushare/dsweb/Get/Document-24400/*.pdf.
Instrument Calibration Requirements (e.g., safety system/device recertification, RF probe calibration)
Access to front end NINO cards for gain matching PMTs, including the top ones via ladder / lift / Gas leak rate check, via access to the gas panel
Inspection Schedules
N/A
References/Associated/Relevant Documentation
Documentation for high voltage and low voltage operation
List of Records Generated (Include Location / Review and Approved procedure)
N/A

Submit Procedure for Review and Approval (See ES&H Manual Chapter 3310 Appendix T1 OSP & TOSP Instructions – Section 4.2 Submit Draft Procedure for Initial Review):
- Convert this document to .pdf
- Open electronic cover sheet:
https://jlabdoc.jlab.org/docushare/dsweb/Get/Document-24048/3310T1Form.doc
- Complete the form
- Upload the pdf document and associated Task Hazard Analysis (also in .pdf format)

Distribution: Copies to Affected Area, Authors, Division Safety Officer
Expiration: Forward to ES&H Document Control

Form Revision Summary / Revision 1.7 – 02/25/2021 – Corrected link to Word doc; updated ‘ESH&Q’ to “ES&H’; other minor edits. No approval required. / Revision 1.6 – 06/23/2020 – Update section 15 to reflect guard number, what to do in an emergency, crew chief numbers, etc. approved by H. Fanning / Revision 1.5 – 04/11/18 – Training section moved from section 5 Authority and Responsibility to section 9 Training / Revision 1.4 – 06/20/16 – Repositioned “Scope of Work” to clarify processes / Qualifying Periodic Review – 02/19/14 – No substantive changes required / Revision 1.3 – 11/27/13 – Added “Owning Organization” to more accurately reflect laboratory operations. / Revision 1.2 – 09/15/12 – Update form to conform to electronic review. / Revision 1.1 – 04/03/12 – Risk Code 0 switched to N to be consistent with 3210 T3 Risk Code Assignment. / Revision 1.0 – 12/01/11 – Added reasoning for OSP to aid in appropriate review determination. / Revision 0.0 – 10/05/09 – Updated to reflect current laboratory operations / ISSUING AUTHORITY FORM TECHNICAL POINT-OF-CONTACT APPROVAL DATE REVIEW DATE REV. / ES&H Division Harry Fanning 04/11/18 02/25/24 1.6 / This document is controlled as an on line file. It may be printed but the print copy is not a controlled document. It is the user’s responsibility to ensure that the document is the same revision as the current on line file. This copy was printed on 8/15/21.