Task Hazard Analysis (THA) Worksheet / (See ES&H Manual Chapter 3210 Appendix T1 / Work Planning, Control, and Authorization Procedure)

Author:	Jack Segal				Date:		08/11/21		Task #: / If applicable
Complete all information. Use as many sheets as necessary
Task Title:		BigBite GRINCH for the SBS experiments			Task Location:			Hall A, downstream of the target and on the beam left side
Division:		Physics			Department:	Hall A		Frequency of use:		Daily from September 2021 onwards
Lead Worker:			Bradley Yale
Mitigation already in place: / Standard Protecting Measures / Work Control Documents				Use of voltage appropriate shielded HV cables and connectors. Restricting orifice and relief valve between regulator and detectors

Sequence of Task Steps	Task Steps/Potential Hazards	Consequence Level	Probability Level	Risk Code (before mitigation)	Proposed Mitigation / (Required for Risk Code >2)	Safety Procedures/ Practices/Controls/Training	Risk Code / (after mitigation
1	Electrical shock from / 2kV high voltage / supplying GRINCH photomultiplier tubes, or / from 6V low voltage power / supply for GRINCH front end / NINO cards	M	L	2	Use of current limited high voltage supply at 2kV. Use of shielded cables and connectors for HV. Use of 4 AWG welding cables and connectors for LV	High voltage cables are only connected or disconnected to/from the detectors, power supplies, and patch panels when power supply is not energized. / ESH Manual Chapter 6200 Electrical Safety Program: / https://www.jlab.org/ehs/ehsmanual/6200.htm / / Training: ESC001, ESC003, ESC004, ESC007, and ESC008 Electrical Safety Awareness Class, Modes, etc. / Electrical safety training for QEW is currently being modified.	1
2	Use of compressed C4F8 gas: / hazards associated with / handling of pressurized / containers and oxygen / displacement in the event of a / catastrophic leak.	M	L	2	Gas supplied through a / pressure regulator / attached to the gas bottle / with flow limited by a / flow meter, an aperature restriction and a pressure relief valve. / Location of the gas bottle in a / non confined place to limit the / effect of oxygen displacement. Each cylinder has 110 pounds of liquid at 25 PSIA. This is 200 cubic feet of gas at STP. The volume of the GRINCH is 82 cubic feet. The area of Hall A is about 21,124 square feet. One cylinder will cover the hall with 0.114 inches of gas. Thus the concern is one of a local ashyxiant. Anytime the detector is opened for repair or maintenance adequate forced ventilation should be used to ensure no stagnant volume of simple asphyxiant remains in the detector. A buddy rule will be instituted for opening the detector after filling.	Setup, connection, or disconnection of / the gas shall only be done by individuals / authorized by this OSP. / ESH Manual Chapter 6150 Gas Cylinder / Safety - Storage, Movement, and / Labeling / https://www.jlab.org/ehs/ehsmanual/ / manual/6150.html, / Training for individuals: SAF130A, / SAF130C / Relief valve and restriction between the regulator and / detector is implemented as a backup in / the event of regulator failure. / Oxygen deficiency hazard SAF103	1
3	Covid-19 contamination (if at an elevated MEDCON level)	M	M	3	Face covering required. / Maintain Social Distancing of 6’ or use appropriate PPE if 6’ distancing cannot be maintained.	OSP: ESH-21-116798-OSP. / Follow required guidelines	1
4	Use of portable ladder / for access to the GRINCH / detector	M	M	3	Use of ladder training and having an observer/assistant.	SAF307 Ladder Safety	1

Highest Risk Code before Mitigation:	3	Highest Risk Code after Mitigation:	1

When completed, if the analysis indicates that the Risk Code before mitigation for any steps is “medium” or higher (RC≥3), then a formal Work Control Document (WCD) is developed for the task. Attach this completed Task Hazard Analysis Worksheet. Have the package reviewed and approved prior to beginning work. (See ES&H Manual Chapter 3310 Operational Safety Procedure Program.)

Form Revision Summary / Revision 0.2 – 07/26/21 – Periodic Review; updated header and footer / Periodic Review – 08/29/18 – No changes per TPOC / Periodic Review – 08/13/15 – No changes per TPOC / Revision 0.1 – 06/19/12 - Triennial Review. Update to format. / Revision 0.0 – 10/05/09 – Written to document current laboratory operational procedure. / ISSUING AUTHORITY TECHNICAL POINT-OF-CONTACT APPROVAL DATE REVIEW DATE REV. / ES&H Division Harry Fanning 08/29/18 07/26/24 0.2 / This document is controlled as an on line file. It may be printed but the print copy is not a controlled document. It is the user’s responsibility to ensure that the document is the same revision as the current on line file. This copy was printed on 8/15/21.