## Signals
|  |  |  | I/O Device |  |  |  |  | I/O Conditioner
Signal description | Range | Tag_name | Accuracy | Qty | Redundant Qty | Description | Price Each | Total Price | Qty | Description | Price Each | Total Price | Expected Accuracy | I/O Signals
Cryogenics
Temperature
Temperature Sensor | > 233K | T_LN_(n)n->corner of the dipole |  | 2 | 0 | PT100 / S65124B foil sensors | 0 | 0 | 2 | Dataforth DSCT-34 | 188 | 376 |  | I: 1X 4 - 20 mA O: 0
Temperature Sensor | > 60 K |  |  | 7 | 6 | PT100/1515 Ceramic | 0 | 0 | 7 | Dataforth DSCT-34 | 188 | 1316 |  | I: 1X 4 - 20 mA O: 0
Temperature Sensor | < 60 K |  |  | 8 | 6 | Thin Film Carbon Glass | 0 | 0 | 8 | Cryomat 1055 Temperature Transmitters | existing |  |  | I: 1X 4 - 20 mA O: 0
VME Rack and Power Supply | N/A | [5 empty] | 0 | 0 | 1 | Lakeshore 2308-1 | 48 | 48 |  | N/A
Gas Flow Controllers
+ Lead He Gas Flow | 0-2000 sccm N2 | V_RCL |  | 1 | 0 | MKS 1749A-02000SV +24Vdc input 4-20mA output | 0 | 0 | 1 |  | existing |  |  | I: 1X 4-20 mA O: 0
- Lead He Gas Flow | 0-2000 sccm N2 | V_LCL |  | 1 | 0 | MKS 1749A-02000SV +24Vdc input 4-20mA output | 0 | 0 | 1 |  | existing |  |  | I: 1X 4-20 mA O: 0
JT LVDT | -6-100% | LVDT_JTV(n) n-> 1-9 |  | 2 | 0 | Existing | 0 | 0 | 2 | Macro Sensors LVC-2402 | 250 | 500 |  | I: 1X 4-20 mA O: 0
JT Valve |  | MOT_JTV(n) n-> 1-6 |  | 2 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 2 | JT Custom Control | 400 | 800 |  | I: 1X Dig O: 2X Dig
Jlab Box (He cool down gas)
JT Valve |  |  |  | 2 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 2 | JT Custom Control | 400 | 800 |  | I: 1X Dig O: 2X Dig
JT LVDT | -6-100% |  |  | 2 | 0 | Existing | 0 | 0 | 2 | Macro Sensors LVC-2402 | 250 | 500 |  | I: 1X 0-10 V O: 0
temperature | 300-73K |  |  | 2 | 0 | PT100 DIN 43760 | 0 | 0 | 2 | Dataforth DSCT-34 | 188 | 376 |  | I: 1X 4 - 20 mA O: 0
Vacuum
LH Thermovac TR206 | Vacuum gauge |  |  | 1 |  | 760 Torr to 1x10-3 Torr |  |  |  | Leybold Combivac CM 31 | existing |  |  | I: 1X 0-10 V O: 0
LH PenningVac PR 35 | Vacuum gauge |  |  | 1 | 0 | 1x10-2 Torr to 1x10-9 Torr | 0 | 0 | 0 | Leybold Combivac CM 31 | existing |  |  | I: 1X 0-10 V O: 0
Pressure
LN2 Pressure | 0.4 - 4 Bar | p_N2 |  | 1 | 0 | RS Components #2493858 | 0 | 0 | 1 | integrated | ? | 0 |  | I: 1X 4-20 mA O: 0
LN2 Switch | 0.4 - 4 Bar | h_LN2 |  | 1 | 0 | Bailey & Mackey Type 14 | 0 | 0 | 0 | Integrated | ? | 0 |  | I: 0 O: 1X Dig
LHe Pressure | 0.4 - 4 Bar | p_He |  | 1 | 0 | RS Components #2493858 | 0 | 0 | 1 | Integrated | ? | 0 |  | I: 1X 4-20 mA O: 0
LHe Switch | 0.4 - 4 Bar | h_Lhe |  | 1 | 0 | Bailey & Mackey Type 14 | 0 | 0 | 0 | Integrated | ? | 0 |  | I: 1X Dig O: 0
Liquid Level
LN2 Level | 0 - 100 % | h_LN2 |  | 1 | 0 | CRYOMAT SARL | 0 | 0 | 1 | Integral to SARL | existing |  |  | I: 1X 4-20 mA O: 0
LHe Level | 0 - 100 % | h_Lhe |  | 1 | 0 | CRYOMAT SARL | 0 | 0 | 1 | Tramsmitter Card | existing |  |  | I: 1X 4-20 mA O: 0
JT Valves
LHe Reservoir | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
LHe Magnet | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
LHe Return | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
15K Cooldown Control | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
LN2 Reservoir | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
LN2 Magnet | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
LN2 Check Valve | -6-100% |  |  | 1 | 0 | Globe Motor (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig O: 2X Dig
Warm Return | -6-100% |  |  | 1 | 0 | Automax (24 VDC @ 0.8A) | 0 | 0 | 1 | JT Custom Control | 400 | 400 |  | I: 1X Dig, 1X 4-20mA O: 2X Dig
Warm Return Temp. | 300-73K |  |  | 1 | 0 | PT100 DIN 43760 | 0 | 0 | 1 | Dataforth DSCT-34 | 188 | 188 |  | I: 1X 4 - 20 mA O: 0
JT LVDTs |  |  |  | 7 | 0 | Existing | 0 | 0 | 7 | Macro Sensors LVC-2402 | 250 | 1750 |  | I: 1X 0-10 V O: 0
Magnet
Voltage Readings
Main Coil Voltages | +/- 300 VDC |  |  | 16 | 16 | AWG-28 wire HABIA cable AB | 0 | 0 | 16 | Dataforth DSCT | 188 | 3008 |  | I: 1X 4-20 mA O: 0
Main Coil Transistion Leads | +/- 300 VDC |  |  | 6 | 0 | AWG-24 wire HABIA cable AB | 0 | 0 | 6 | Dataforth DSCT | 188 | 1128 |  | I: 1X 4-20 mA O: 0
Power Supply Voltage | +/- 10 VDC |  |  | 1 |  | N/A | 0 | 0 | 1 | Dataforth DSCT | 188 | 188 |  | I: 1X 4-20 mA O: 0
Strain Gauges
Gauge | 0 kN to 35 kN | F_[L/R(n)] |  | 8 | 8 | CEA-06-125UT-350 Micro-measurement | 0 | 0 | 8 | Dataforth SCM5B-38-33 | 190 | 1520 |  | I: 1X -5-5 V O: 0
Mounting Plane | [5 empty] | N/A | 0 | 0 | 1 | Dataforth SCMPB07-02 | 247 | 247 |  | N/A
PLC commuications |  |  |  | 1 | 0 | Metrolab Field Regulation RG2040 | [7 empty] | RS232
Power supply Communications |  |  |  | 1 | 0 | Metrolab Field Regulation RG2040 | [7 empty] | RS232
Analog adjust | +/- 20 mA |  |  | 1 | 0 | Metrolab Field Regulation RG2040 | [7 empty] | +/- 20 mA
Power Supply
Dump System
Reversing Switch
Position Indicator |  |  |  | 1 | 0 | Short = OK | [7 empty] | I: 1X Dig O: 0
Position OK |  |  |  | 1 | 0 | Short = Positive | [7 empty] | I: 1X Dig O: 0
Keep-Alive
Reset |  |  |  | 1 | 0 | 24VDC = Restart Timer | [7 empty] | I: 0 O: 1X Dig
Status |  |  |  | 1 | 0 | Short = OK | [7 empty] | I: 1X Dig O: 0
Quench Detector
Main Coils |  | MPS_Quench |  | 2 | 0 |  |  |  | 1 | DANFYSIK 4 Channel
Main Coils Transistion Leads |  |  |  | 2 | 0 |  |  |  | 0
Total Price | [7 empty] | 0 |  |  |  | 15945
[10 empty] | with Overhead and Contingency @ 50% |  | 23917.5
Notes:

## IO
Signal |  | x3
Inputs
4-20 mA | 56 | 168
+/- 10 VDC | 24 | 72
24 VDC | 24 | 72
Outputs
24 VDC | 32 | 96