## Slide 1: Danfysik PSU Fast Dump Opening Timing
Steven Lassiter
Eric Sun
Mike Fowler

## Slide 2: Danfysik system 8500 PSU

## Slide 3: Dump Switch Response
Power Supplies dump switch opening times were much slower than expected, of the order of 750 to 900ms.
Transistor bank V-Loop driver appeared to be partially tuning back on and then off during the mechanical opening of the dump switch, due to shock vibrations possible.

## Slide 4: Dump Switch

## Slide 5: Initial Fast Discharge Timing HB
Chart: 

## Slide 6: Modifications
Moved dump switch’s holding coil’s drain resistor, RE, from AC supply side to DC supply of rectifier.
Increased drain resistor, RE, from 1M Ω to 2.5M Ω.
Added alternative interlock signal path directly to disable PSU’s transistor bank voltage-loop driver. Bypasses chain of mechanical relays.

## Slide 7: Moving drain resistor “Re”
AC Power to Dump Coil
DC Power to Dump Coil

## Slide 8: 400A FD
Chart: Time

## Slide 9: Conclusions
Dump Switch opening time has been improved from >750ms to 37ms.
PSU starts to ramp current down as soon as 16ms after receiving an interlock.
Relay bouncing due to mechanical shock has been eliminated by adding a parallel interlock to disable the V-loop driver sooner.