# E P chamber bleed up checklist		Date __/__/__

# Revised  12/16/98

Name of persons doing checklist__________________________________________

### NOTE: ITEMS IN RED REPRESENT A POSSIBLE HAZARD TO EQUIPMENT

Prior to beginning this checklist you should

1: Record the vacuum in the EP chamber in Beam line vacuum rack _____________

2: Record the vacuum in the Moeller Beam line in Beam line vacuum rack ___________

3: Record the vacuum in the Target chamber in Beam line vacuum rack _____________

4: Record the vacuum in the Exit beam pipe in Beam line vacuum rack _____________

5: Record the Vacuum in the Compton Beam line in Beam line vacuum rack __________

## To bleed up EP vacuum chamber

___	Call MCC and inform them that you are taking the Hall A beam line valves to the maintenance mode in order to bleed up the E P chamber. Ask them to enter this in the E-log, and be sure to get the controller’s name that took the call.

operator’s name ____________________________

___ 	Insure that all of the manual switches on the Hall A beam line Vacuum Valve control chassis located in Beam line vacuum rack _____________	 are in the closed position

___ 	Turn the key to maintenance, then remove and tag the key.

___	Visually verify that all Hall A beam line valves are closed

___	Close the gate valve between the EP chamber and it’s turbo pump

___	Turn off turbo pump

___	Turn off backing pump and automatic valve

___	If only shutting EP system down stop here

___	Using dry N2 bleed up EP chamber through K F-16 VAT valve being careful not to over pressurize the chamber

## To pump down EP vacuum Chamber

___	Turn on backing pump and automatic valve

___	Open vacuum valve between EP and roughing pump

___	Pump EP chamber down to < 10 mtorr

___	Verify that turbo is functioning normally

___	Open EP to turbo isolation valve

___	Close vacuum valve between EP and roughing pump

___	allow vacuum in the EP chamber to reach 5x10-4

___	verify that vacuum in the Moeller beam line is< 1x10-4

___	Using the manual switch, open the beam line valve between the EP chamber and the Moeller turbo.

___	Using the manual switch, open the beam line valve between the EP chamber and the raster.

___	verify that vacuum in the E P chamber is ~1x10-4

___	Using the manual switch, open the beam line valve between the EP chamber and the downstream end of the differential pumping station.

___     If the Cryo target is installed, verify that the target chamber vacuum is <1x10-4

___	If the Cryo target is installed. using the manual switch, open the beam line valves on both ends of the target chamber.

___	Recheck all of the Hall A vacuum systems and insure that all beam line vacuums are at least 1 x 10-4.

___	Call MCC and inform them that you are taking the Hall A beam line valves back to the operational mode. Ask them to enter this in the E-log, and be sure to get the controller’s name that took the call.

operator’s name ____________________________

___	 Untag and insert the key, then turn it to the operational position

___	Visually verify that all Hall A beam line valves are open