Cryo Target Pre Beam Checklist

Last revised 9/20/04 	Date______time _______

#### This checklist will be performed after every restricted access to Hall A that maintenance is performed

Person responsible for checklist __________________________________

People checking list _______________________________________________________

###### Left Arm

##### Spectrometers

___	Correct angle ___________ (not to be used for calculations)

___	Correct pointing _________(not to be used for calculations)

___	Collimator operation at 3 positions

___	check camera angle for movement in both directions

___	Bogie controls checked for operation (do not move)

___	Bogie power is 	ON ___ 	Off ___

___	Check spectrometer for obstructions to movement

___	Check intergen bottles for correct pressure

___	Ensure that Intergen alarm switch is in the normal position

___	Insure that 14-degree stop pin is installed (if used)

___	Insure that outer limit stop is installed (if used)

___	minimum/Maximum angles for spectrometer from _____Deg. To ____ Deg.

___	magnetic shielding installed (if necessary)

___	pointing surveys complete

Vacuum

___	Insure that turbo backing pump is on

___	Turbo on at turbo controller in rack # 1H71B01

___	Pump valves open at valve controller in rack # 1H71B01  channel #2

___	Convectron gages read “0” millitorr

___	ensure roughing pump at dipole entrance has sufficient oil

___	Cold cathode gages on at gage in rack # 1H71B01 cold cathode < 5x10-5

___	Actual cold cathode reading ________

___	Entrance & exit vacuum windows functional

___	ensure that the Q3 insulating turbo and backing pump are on and functional

___	ensure that the automatic valve is open and that the Convectron gage reads 0

Left Arm

Magnet controls

Q1

___	Q1 full of liquid  (80%) actual______

___	Open lead flows on Q1 to 80 slm as read from rack #Q171Q

___	Actual lead flows	A______	B______

___	Cctv camera on and focused

D1

___	Dipole full of liquid  (80%) actual______

___	Open lead flows on Dipole to 80 slm as read from rack #D171Q

___	Actual lead flows	A______	B______

Q2

___	Q2 full of liquid  (80%) actual______

___	Open lead flows on Q2 to 60 slm as read from the Q2 instrument rack meter.

___	Actual lead flows	A______	B______

___	Insure that lead heaters are on and operating

Q3

___	Q3 full of liquid  (80%) actual______

___	Open lead flows on Q3 to 60 slm as read from the Q3 instrument rack meter.

___	Actual lead flows	A______	B______

___	Insure that lead heaters are on and operating

Left Arm

Power supplies

POWER SUPPLY TURN ON PROCEDURES

___	Verify UPSs as operational on all power supply controls (with no current on magnets only)

___	Red rotating beacons on

Q1:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Visually check power supply for faults

___	When all faults have been cleared, insure that power supply is in remote control

Q2:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on both sets of three pole breakers located on power supply

___	Visually check power supply for faults

___	When all faults have been cleared, lift lever on lower right side of supply

___	Insure that power supply is in remote control

Q3:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on both sets of three pole breakers located on power supply

___	Visually check power supply for faults

___	When all faults have been cleared, lift lever on lower right side of supply

___	Insure that power supply is in remote control

Dipole:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on power lever on right upper side of supply

___	Visually check power supply for faults on supply and at rack #

___	When all faults have been cleared, insure that power supply is in remote control

___	Cctv camera on and focused

___	Check power supply for proper polarity 	positive____	negative____

___	NMR gradient compensation for on and proper polarity

positive____ negative____

###### Right Arm

##### Spectrometers

___	Correct angle ___________ (not to be used for calculations)

___	Correct pointing _________(not to be used for calculations)

___	Collimator operation at 3 positions

___	check camera angle for movement in both directions

___	Check spectrometer for obstructions to movement

___	Check intergen bottles for correct pressure

___	Ensure that Intergen alarm switch is in the normal position

___	Insure that 15-degree stop pin is installed

___	Maximum/minimum angles for spectrometer _______to_______

___	magnetic shielding installed (if necessary)

Vacuum

___	Insure that turbo backing pump is on

___	Turbo on at turbo controller in rack # 1H71B01

___	Pump valves open at valve controller in rack # 1H71B01  channel #2

___	Convectron gages read “0” millitorr

___	ensure roughing pump at dipole entrance has sufficient oil

___	Cold cathode gages on at gage in rack # 1H71B01 cold cathode < 5x10-5

___	Actual cold cathode reading ________

___	Check dipole turbo for on

___	ensure that the dipole insulating turbo and backing pump are on and functional

___	ensure that the automatic valve is open

Right Arm

Magnet controls

Q1

___	Q1 full of liquid  (80%) actual______

___	Open lead flows on Q1 to 80 slm as read from rack #Q171Q

___	Actual lead flows	A______	B______

___	Cctv camera on and focused

D1

___	Dipole full of liquid  (80%) actual______

___	Open lead flows on Dipole to 80 slm as read from rack #D171Q

___	Actual lead flows	A______	B______

Q2

___	Q2 full of liquid  (80%) actual______

___	Open lead flows on Q2 to 60 slm as read from the Q2 instrument rack meter.

___	Actual lead flows	A______	B______

___	Insure that lead heaters are on and operating

Q3

___	Q3 full of liquid  (80%) actual______

___	Open lead flows on Q3 to 60 slm as read from the Q3 instrument rack meter.

___	Actual lead flows	A______	B______

___	Insure that lead heaters are on and operating

###### Right Arm

Power supplies

POWER SUPPLY TURN ON PROCEDURES

___	Verify UPSs as operational on all power supply controls (with no current on magnets only)

___	Red rotating beacons on

Q1:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Visually check power supply for faults.

___	When all faults have been cleared, insure that power supply is in remote control.

Q2:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on both sets of three pole breakers located on power supply

___	Visually check power supply for faults

___	When all faults have been cleared, lift lever on lower right side of supply.

___	Insure that power supply is in remote control

Q3:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on both sets of three pole breakers located on power supply

___	Visually check power supply for faults.

___	When all faults have been cleared, lift lever on lower right side of supply.

___	Insure that power supply is in remote control

Dipole:

___	Visual inspection of main current leads, dump resistor, and lead flags (for condition, visual shorts, etc.)

___	Unlock power disconnect switch and turn on AC power

___	Turn on power lever on right upper side of supply.

___	Visually, check power supply for faults on supply and at rack #OD172Q.

___	When all faults have been cleared, insure that power supply is in remote control.

___	Cctv camera on and focused

___	Check power supply for proper polarity 	positive____	negative____

___	NMR gradient compensation for on and proper polarity

positive____	negative____

Target

___	Windows functional

___	Cctv cameras “on” and focused

___	Target light “on”

___	Backing pump “on” at pump

___	ensure roughing is closed

___	Turbo “on” at rack # 1H75B09

___	Turbo valve “open” at rack # 1H75B09 channel # 1

___	ensure target convectron set point is 5 torr

___	Gages operational

___	Convectron “0” millitorr at rack # 1H75B09

___	Cold cathode < 5x10-4 at rack # 1H75B08

___	Actual cold cathode reading __________

Exit beam tube

___	Diffuser cooler on

___	Diffuser water level ok

___ 	close flow valve and observe flow meter (drops to 0)

___	open flow valve and observe flow meter (rises to 1 gpm)

___	Backing pump is “on” and operational

___	Valve “open” at pump

___	Turbo “on” at rack # 1H75B09

___	Gages operational

___	Convectron “<5” millitorr at rack # 1H75B09

___	Actual convectron gage reading ________

___	magnetic shielding installed (if necessary)

Entrance beam tube

___	Insure that beam line girder turbo is on and running

___	Insure that there is cooling water flow to the Moeller Dipole

___	Insure that E P turbo is on and running

___	Instrument air compressor functioning normally

___	Beam line vacuum valves “open” (visually checked)

___	Call MCC, get the name of the person you talked to and say “I am doing the Hall A pre beam checklist, Please Insure that the Hall A beam line valves are set to close” after they say that they are, say “I am turning the control key from maintenance to operational are you ready” after they say yes, turn key and tell them “you have control could you please open the valves so that we can verify operability make an e-log entry”

Controls check from the computer console

___ 	Pull up the Hall A tools page

___ 	Ensure that all of the lead flows are in the green

___ 	Ensure that all liquid levels are in the green

___ 	Ensure that all polarities are correct

___ 	Using the current button open the control page to left Q1

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to left Q2

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to left Q3

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to left D1

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to right Q1

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to right Q2

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to right Q3

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___ 	Using the current button open the control page to right D1

___ 	Clear all faults and turn on magnet with correct polarity

___ 	Ramp magnet to 100 A

___	input .5 GeV/c for both spectrometers

___	ensure that all magnets lock in for the input momentum

___	 list magnets that do not ________________________

## Hall

___	All interlocks in rack # 1H75B08 indicate green

___	Check 3 Moeller power supplies for on and in remote

___	Ensure installation of Ion chambers at EP, target, and 2 at beam dump

___	Correct LCW flow and pressure (100 psi supply and 60 psi return)

___	Cctv cameras on and focused

___	Cctv monitors at X terminal off

___	Clear of unnecessary equipment

___	Man lift and Forklift in truck ramp.

___	ensure that all lifting slings and safety harnesses are correctly stored and that the storage cage is at least 90 deg from the beam dump and at least 60 ft from the target

___	Perform pre sweep of run safe boxes.

___	Unnecessary personnel exit Hall.

___	Move Electron spectrometer stairs clear of lower balcony.

___	Ensure polar crane is positioned over the entrance beam pipe,

and that power is off at the power disconnect switch

___	Ensure that spectrometer entrance window guards are removed

___	Ensure that spectrometer exit window guards are removed

___	Ensure that detector VDC covers are removed

___	Ensure that target window guards are removed

___	Deliver checklist to shift leader or run coordinator

Name of person checklist was delivered to ______________________