#### User name S. Strauch

#### Log entry time 12:27:01 on September 30, 2006

#### Entry number 185963

keyword=HRS magnet cryo-level problems

I discussed the cryo-level problem of the HRS magnets with Mark Stapleton. The problem is not yet fully understood. It happened this morning even after Mark increased the cryo levels last night following the previous problem. In his recent CLOG entry he states: "While looking at the system, it appears the magnets have quenched. The only thing wrong with that idea is the flow to recovery did not change in the manor of a 4 magnet lose of helium?"; see also the other CLOG entries about this issue.

A copy of this log entry has been emailed to: folts@jlab.org

#### User name Fansler

#### Log entry time 10:21:29 on September21,2006

#### Entry number 184336

keyword=Left Q2 Ramp Problem

Left Q2 would not cycle above 1570 amps. Fault indications on GUI were "Quench" and "Interlock". After reviewing the stip charts, we found no evidence of an actual quench. Hall C is currently cooling their target, we believe it is effecting our lead flows. Right Q3, and Left Q3 tripped with the same indications. I attempted to cycle left Q2 with the same results. I increased the lead flow by 5 percent, and Q2 cycled normally. Both Q3's were reset remotely, and cycled normally. All magnets are operational at this time. It is likely we will see more magnet trips while the hall C target cooldown continues.