#### User name fansler

#### Log entry time 08:32:13 on February 09, 2007

#### Entry number 191681

keyword=R_Q2_Liquid _Level_Trip

0800 I noticed the right Q2 was tripped on an interlock fault. Upon further investigation, I discovered the liquid level was oscillating between 7 percent and 90 percent. I called Pete, and requested he tune the PID loop to overcome the oscillations. I reset the magnet and called the counting house to request they reset current and ramp the magnet.

A copy of this log entry has been emailed to: fansler@jlab.org