#### Log entry time 12:06:05 on October 22, 2005

#### Entry number 157681

keyword=tech/right spectrometer quenches

This morning we had trips on the right Quads. I believe that this is because the right dipole is overfilling. the local readout reads 100% and the GUI reads 71% this discrepency needs to be resolved.

The top fill valve is being controled to full open. I believe that it may be the valve card and cycle power, the valve still drives closed.

Called cryo tech on call, cycled power on LL meter, valve still drives closed.

Decided to change valve card but when I got back with the card the system was working properly. Karma? Future TOCs should be aware of this problem as I suspect the card is dieing from radiation.

A copy of this log entry has been emailed to: fansler@jlab.org