#### User name folts

#### Log entry time 14:37:00 on August 25, 2005

#### Entry number 153281

keyword=techs/Q3

Q3 has been quenching due to liquid level for the last 2 days

We have been running a diagnostic strip chart to understand the problem

All indications were that the insulating vacuum had deteriorated, but before we could get an access to verify that this was the only problem the valves were propped open more by cryo, and then we lost the cryo IOC to the right arm.

When we finally got an access we pumped Q3 anyway (the vacuum gauge read 0 motor) for 2 hours. We seem to be using a lot of ESR capacity and thought this might help.

Called J. Creel for help on the IOC problem.

A copy of this log entry has been emailed to: fansler@jlab.org

Problem was liquid level meter.  We looked at a strip chart of upper and lower valves.  They showed 100% and no movement for several hours.  Replaced liquid level meter, and resolved the problem.