#### User name S. Strauch

#### Log entry time 12:27:01 on September 30, 2006

#### Entry number 185963

keyword=HRS magnet cryo-level problems

I discussed the cryo-level problem of the HRS magnets with Mark Stapleton. The problem is not yet fully understood. It happened this morning even after Mark increased the cryo levels last night following the previous problem. In his recent CLOG entry he states: "While looking at the system, it appears the magnets have quenched. The only thing wrong with that idea is the flow to recovery did not change in the manor of a 4 magnet lose of helium?"; see also the other CLOG entries about this issue.

Solution:

The hall 4k supply valve was not opened to it’s normal position, thus reducing the overall flow to the hall.  Problem was detected and corrected by cryo group