1. Define the scope of work

- Job Summery

Pressure test the Moller target reservoir to 30 psi gauge

- Affected systems

Moller target reservoir

2. Analyze Hazard

Task Hazard Analysis Worksheet

| Complete all information. Use as many sheets as necessary. |  |
| --- | --- |
| Task location PSB | Task title Moller target N2 shield pressure test |
| Division Physics | Prepared by Folts |
| Department Hall-A | Reviewed by Dezern |
| Supervisor Folts | Approved by Folts |
| Standard Requirements: (PPE, TOSPs, SOPs, etc.) |  |

| Task | Potential Hazards | Risk Code before mitigation | Safe Procedures/Practices/Controls | Risk Code after mitigation |
| --- | --- | --- | --- | --- |
| Moller target N2 shield pressure test | Ear damage, Eye damage | 2 2 | Ear protection Safety glasses Test preformed in controlled area Stay at lease 3 feet from vessel while under pressure | 0 0 |

3. Develop and implement Hazard controls

- List of PPE

Ear plugs

Safety glasses

- Identify danger zones

Minimum 3 ft. distance from reservoir

- Training required

This procedure

4. Perform work within controls

- Number of people needed

2

- List of materials

Reservoir

Calibrated pressure gauge

- Procedure

Set up calibrated pressure gauge in series with the reservoir

Cap off and secure the 2 remaining openings to the reservoir

Secure the test area

Slowly pressurize the reservoir to 30 psi gauge.

Hold pressure for 10 minutes and then visually check the reservoir for deformation.

- Post job briefing and cleanup

De-pressurize and remove instrumentation

5. Provide feedback and continuous improvement.