Hall A spectrometer LVDT reset procedure

By:  Scot Spiegel

This procedure is to be performed by qualified Hall “A” personnel only.

If the numerical readout of any LVDT becomes corrupted and cannot be reset from the X-terminal this procedure will be used.

The following is an outline of the steps needed to RESET the LVDT communications controls and controller for any of the LVDT’s utilized for spectrometer alignment.

Begin by re-establishing communications with the LVDT controller.

Step 1.  Pull up the MEDM screen from the X-terminal and “click on” File.

Open the file list. There will be a list of files on the left selection box.

Step 2.  Scroll down the directory files until you see “disp” and double click on it.

This gives you the “Display” directory.

Step 3.  Scroll down the display directory until you see one of the following:

a. “lisp/HRSL” This is the file for the LEFT spectrometer alignment OR

b. “lisp/HRSR” This is the file for the RIGHT spectrometer alignment.

“Double click” the file selection for the spectrometer requiring a reset.

Step 4.  The file list that appears to the right of the directory are files for the

user screens displayed on the X-terminal. The particular LVDT can be

selected from this screen.

Step 5.  Scroll down the list until you can select the correct file for the LVDT

requiring communications re-establishment. The files will be listed as

follows:  1. “lisp/HRSL”   OR   2. “lisp/HRSR” (see below)

- “HacL_HPM_LVDT.adl” This is the file for the horizontal position of

the LEFT spectrometer.

- “HacL_VPM_LVDT.adl” This is the file for the vertical position of

the LEFT spectrometer.

2. “lisp/HRSR”

- “HacR_HPM_LVDT.adl” This is the file for the horizontal position of

the RIGHT spectrometer.

- “HacR_VPM_LVDT.adl” This is the file for the vertical position of

the RIGHT spectrometer.

“Double click” on the appropriate user screen file to pull it up.

Step 6.   Verify the parameters of the LVDT controller from the list in the center

of the screen. They should appear as follows:

A-Mult	5

A-Fsval	22500

A-Zero	0

A-Calib	0

A-Slope	1.000000

A-Prest	0

Hys-hSP	0

Hys-lSP	0

Display			chA

Step 7.   If the parameters do not match change them in the user fields to the right

of the display list.

Step 8.   Press red RESET button to re-establish LVDT controller

communications.

Once communication is established the LVDT controller will need to be reset

Step 9.   Recycle power to appropriate LVDT controller by disconnecting AC

power cord from back of the unit.  They are located on the magnet

power supply platform in the control rack.

Step 10. Check X-terminal for new alignment reading. It should be a value

containing no more than 5 digits.