Right Dipole Octagon computer flying reset procedure/checklist

By James Proffitt and Mark Stevens

Revision:  November 1, 2000

Persons performing procedure: ________________________________________________date________

This procedure/checklist outlines the flying reset/restart of the Right Dipole power supply Octagon computer, the Dynapower power supply’s imbedded controller.  This procedure is to be used when the Right Dipole Dynapower no longer responds to EPICS control inputs, and the Octagon computer has exited its program, but is otherwise operational.  It is intended to be performed by qualified Jefferson Lab technicians who have been trained to perform it by the Hall A technical staff.

Note:

The Right Dipole is not identical to the Left Dipole.  The Right Dipole has an internal short circuit in the magnet copper windings.  This gives the system some unique operational constraints, and as a result, a unique protection system and operational software.  Both the external (EPICS) and internal (Octagon) software are unique to this system.   A copy of the Octagon software is held on file by the Hall A technical staff.

# Caution

The Right Dipole is a low-voltage, superconductive magnet system capable of storing tremendous amounts of energy.  Because of the shorted windings of the magnet, the power supply must not be arbitrarily turned off.  Every effort must be made to ramp the current down to 0 amps prior to turning the power supply off.  Sudden excursions of the current, such as turning the power supply off, can result in a quench of the magnet, with most of the energy stored being released inside the magnet.  Severe equipment damage could result.

1. _____  Take the black laptop computer (the Dell Latitude LX), a DB-9 to RJ-11 adapter and a short (6’) RJ-11

(telephone type) cable to the Right Spectrometer power supply platform.

2. _____  In rack 1H75B05 disconnect the RJ-11 cable marked “0” from the front of the serial cross connect chassis.

It is the most right hand connector as you look at the front of the chassis.

3. _____  Boot the laptop.  Under program manager “accessories” icon, click “terminal”.  In terminal mode,  pull

down the “file” menu.  Under the “file” menu, click  “open”.  In the “open” window, select “hdynapwr”

and then click “OK”.

4. _____  Connect the laptop serial port to the RJ-11 receptacle opened in step 3 above, using the DB-9 to RJ-11

adapter and the 6’ RJ-11 cable.

5. _____  On the laptop, press “return”.  A carat ( > ) cursor should appear at the left margin of the page.  This

tells you that the processor is operating, and you are communicating with it.

6. _____  Type “list”.  The program should print to the screen, beginning with:

5 . . Program File Name is 32131WNG

6 . . August 19, 1993

7 . .

8 . . Modified 10-Mar.-1998 by James Proffitt, Jefferson Lab

(etc.)

7. _____  Type the following lines:

DEL 0- (return)

LOAD (return)

DEL 1260 (return)

DEL 1290-1350 (return)

30 Amps = (the value read from the EPICS control screen current readback from the Dynapower)(return)

31 CO = (the value read from the EPICS control screen current readback from the Dynapower)(return)

RUN (return)

8. _____  The Dynapower is now operating on the newly modified software.  Disconnect the laptop from the

Dynapower at the RJ-11 connector in rack 1H75B05.

9. _____  Reconnect the cable marked “0”.

10. ____  In the EPICS control window for the Right Spectrometer Dynapower, enter a current set value of  0  amps.

The power supply should now begin to ramp to  0  amps.

11. ____  After the power supply has reached  0  amps, the Octagon computer’s original software must be restored.

# Caution

The Dynapower power supply is operating with the modified software.  Do not input any other commands except for step #10, above.  The Octagon normal operating software must be restored in this step before normal operations can resume.

Turn off the main input power switch on the front of the Dynapower power supply, PD172, a large lever-         type switch with a red and black handle on the right side of the front panel. Unplug the 120VAC plug which provides uninterruptible power to the Dynapower.  This plug is located inside rack OD172Q, is plugged into a rack-mount uninterruptible power supply, and is marked by a CEBAF caution tag “CAUTION – Uninterruptible control power for the Hadron Dipole power supply – do not disconnect”.

Wait 30 seconds to let the various internal voltages to subside.  Then reapply power by first plugging in

the 120VAC plug and then turning on the main input power switch.  The Octagon computer autoboots

from a PROM device.

12. ____ Using Hall A EPICS software, reset faults and test for proper operation.