RIGHT DIPOLE QUENCH DETECT

NEWPORT CONTROLLER

SCALING & PROGRAMMING INSTRUCTIONS

INFCP-011

ANALOG OUT

The menu button will allow you to scroll through all the menu headings in the controller.  The tare button will allow you to enter the submenus.  Once in the submenus the Max button will allow you to increase through the options for each setting, the tare button will allow you to move from left to right through the display.  If you wish to change a setting scroll through until you see your choice displayed on the  screen, then press the tare button to enter the submenu.  Use the max button to increase digits, or scroll through options.  Use the tare button to move to the right through the digits.  The reset button performs the same function as the back button on you PC desk top (if you only press it once). Pressing the reset button twice will reset the controller and take you back to the static display.  After making any changes press menu to store.  The controller will then display the next menu item.  Below are the settings that should be programmed into the right dipole magnet quench detect controller.  Be sure to check the set point references for each individual controller, they may differ.   You must perform the scaling procedure on each newly installed controller.

SETPOINT 1			1.2 VOLTS		ACTIVE BELOW

SETPOINT 2			-1.2 VOLTS		ACTIVE ABOVE

SCALING (INFCP-011B)

1.	Apply 0V  to input

2.	Press the menu key until R.d.s.o. is displayed	(Reading scale and offset)

3.	Press tare

IN 1							(Input one)

4.	Press tare

Last stored value

5.	Press tare

Actual input

6.	Press menu to store

Rd1							(Read one)

7.	Press tare

Last stored value

8.	Press max button to change display to 0.00

9.	Press menu

IN2							(Input one)

10.	Apply 1 V to the input

11.	Press tare 2 times

12.	Press menu

RD2							(Read one)

13.	Press tare

Change the display to read 1.00

14.	Press menu

PROGRAMMING (INFCP-011) (0-5v)

1.	Press menu

INPt					(Input)

2.	Press TARE

+/- 5V					(Input signal range -5V to +5V)

3.	Press menu

dEC.P					(Decimal point)

4.	Press tare

F.FFF					(Decimal point placement)

5.	Press menu

Rd.S.O.				(Reading scale and offset)

6.	Press tare

IN  1					(Input one)

7.	Press tare

-.796	(Minimum input signal / or last stored value for input 1)

8.	Press menu

Rd  1					(Read one)

9.	Press tare

-1.999	(Minimum input display value / last stored value for read 1)

10.	Press menu

IN  2					(Input two)

11.	Press tare

.796	(Maximum input signal / last stored value for input 1)

12.	Press menu

Rd   2					(Read two)

13.	Press tare

1.999	(Maximum input display value / last stored value for read 1)

14.	Press menu

Rd.CF					(Reading configuration)

15.	Press tare

R.1=T					(Tare button enabled)

16.	Press tare

R.2=4	(Input resolution - 10uV for unipolar units, 25uV for bipolar units…see page 29 of operators manual)

17.	Press tare

R.3=F					(Filtered value displayed)

18.	Press menu

S1.CF					(Set point one configuration)

19.	Press tare

S.1=B					(Set point one is active below the set point)

20.	Press tare

S.2=U					(Set point one is unlatched)

21.	Press menu

S2.CF					(Set point two configuration)

22.	Press tare

S.1=A					(Set point two is active above the set point)

23.	Press tare

S.2=U					(Set point two is unlatched)

24.	Press menu

S1.db					(Set point one dead-band ) (not active)

25.	Press tare

.003					(hysteresis level of set point one, if active)

26.	Press menu

S2.db					(Set point two dead-band ) (not active)

27.	Press tare

.003					(hysteresis level of set point two, if active)

28.	Press menu

Ot.CF					(Output configuration)

29.	Press tare

0.1=E					(Analog output enabled)

30.	Press tare

0.2=V					(Analog output = voltage)

31.	Press tare

0.3=A					(Analog output = regular)

32.	Press menu

Ot.S.O					(Output scale and offset)

33.	Press tare

Rd1					(Read one)

34.	Press tare

0.000	(Minimum output display value stored during scaling, or last stored value for read one)

35.	Press menu

OUt.1					(Output one)

36.	Press tare

00.00	(Minimum output signal stored during scaling, or last stored value for output one)

37.	Press menu

Rd  2					(Read two)

38.	Press tare

1.000	(Maximum output display value stored during scaling, or last stored value for read two)

39.	Press menu

OUt.2					(Output two)

40.	Press tare

01.00	(Maximum output signal stored during scaling, or last stored value for output one)

41.	Press menu

Lk.CF					(Lockout configuration)

42.	Press tare

SP.=E					(Set point changes enabled)

43.	Press tare

RS.=E					(Reset button enabled in run mode)

44.	Press tare

SP.=E					(Set point changes enabled)

45.	Press set point

SP1					(Set point one)

1.2					(Set point one value)

47.	Press set point

SP2					(Set point two)

-1.2					(Set point two value)

Controller should revert to static display

-.127 to .95