- Summary of job

This procedure is to de-sting the left Q1 u-tube, sting the “Y” can, and re-sting the left Q1 u-tube.

- Affected systems

Left/Right spectrometer cryogenic systems

Left/Right spectrometer magnet power supplies systems

ESR

- List of PPE

Safety glasses

Gloves

Long sleeves

Hard hat

Harness

Ladder

Man lift

- Identify danger zones

Areas located directly above U-tube/”Y” can removal/installation

Areas located below the u-tube/”Y” can in a 45 degree cone

- Number of people needed

4 techs to handle u-tube and “Y” can

1 crane operator

Cryo support

- List of materials

“Y” can

Snoop

Strap wrench

Acetone and rags

Vinyl   tape

Purge line (N2 and He)

Crane

Die (thread chaser)

Heat Gun

- Procedure

Pre job briefing consisting of assignments, sequence and a review of hazards

Pre-sting

Ensure the house 4 atmosphere He is on for the purge

Extend ladder to Q1/Q2 service can.  Tie off ladder.

Extend He purge line to service can platform

De-sting Q1 He supply U-Tube

Ensure Q1 cold return valve and lead flow valves are open

Cryo shut off 4.5k supply flow to the spectrometer

Check the local pressure gauges on the service cans and the magnet for a suitable low reading (a few PSI)

Attach crane to installed u-tube

Loosen bayonet nuts on installed u-tube

Pull upward on the u-tube both ends		2 techs

Raise the u-tube with the crane	1 tech

Close ball valves immediately after the bayonets clear the valves (magnet and service can end)		2 techs

Chase threads on service can bayonets

Check bayonet socket for debris and moisture

Place u-tube on floor

One tech begin Q1 U-tube pre-sting section of procedure

Pre-sting “Y” Can

At least 24 hours prior to sting begin dry N2 purge on “Y” can

- Open both female bayonet ball valves

- Install N2 purge line in male bayonet

- Ensure flow is exiting both female bayonets

- Pull vacuum on “Y” can at least 24 hrs prior to sting, ensure o mtorr

- Check u-tube o-rings

- Chase threads on “Y” can bayonets

- Check bayonet socket for debris and moisture

Sting  “Y” Can

Remove N2 purge, Install dummy bayonets in both female bayonets

Crane “Y” can into position, place “Y” can u-tube on the closed He supply bayonet

Remove vinyl tape and clean bayonet ends with acetone

Begin He purge using the installed valve located on the service can

Purge for three minutes with house He

Ensure Q1 cold return valve and lead flow valves are open

Ensure Cryo shut off 4.5k supply flow to the spectrometer

Check the local pressure gauges on the service cans and the magnet for a suitable low reading (a few PSI)

Cease purge

Install u-tube

- Lower crane 			(1 tech)

- Open ball valve		(1 tech)

- Push u-tube into bayonet	(1 tech)

- Tighten bayonet nut and snoop

Pre-sting Q1 u-tube

Defrost U-tube

Install N2 purge

Clean u-tube bayonets

Check u-tube o-rings

Sting Q1 U-Tube

Remove N2 purge, tape both bayonet ends

Crane Q1 u-tube to service can platform

Remove vinyl tape and clean bayonet ends with acetone

Begin He purge using the installed valve located on the service can.  Purge for three minutes

Ensure Q1 cold return valve and lead flow valves are open

Ensure Cryo shut off 4.5k supply flow to the spectrometer

Check the local pressure gauges on the service cans and the magnet for a suitable low reading (a few PSI)

Cease purge

Install u-tube

- Lower crane 		(1 tech)

- Open ball valves	(2 techs)

- Push u-tube into bayonet’s	(2 techs)

- Tighten bayonet nuts and snoop

Re-initiate 4.5K supply flow

- Post job briefing and cleanup

Examine the task for hazards

Frostbite

ODH

Fall

Crane

Eye

Risk codes for hazards

Frostbite	2

ODH		1

Fall		2

Crane		1

Eye		1

Mitigate hazards

Frostbite	gloves

ODH		training

Fall		harness

Crane		training

Eye		glasses, clean area

Risk codes after mitigation

Frostbite	1

ODH		1

Fall		1

Crane		1

Eye		0