#### User name folts

#### Log entry time 13:34:11 on October 24, 2005

#### Entry number 157872

keyword=tech/right dipole

Began troubleshoot of 04 valve full open with cryo, valve checkd good.

Heidi sugested check of LL meter, noted that the 30% offset was still there. Control point for the LL at cryo was 71%. with the offset all we could get to was 70.1%.

Reset cryo set point to 68% and the valve closes, problem seems to be in the meter or the isolation module.

During process the right dipole quenched, awaiting refill......

A copy of this log entry has been emailed to: fansler@jlab.org