#### User name folts

#### Log entry time 16:44:47 on August 15, 2005

#### Entry number 152060

keyword=tech/ magnet quench

according to Pete at cryo the lead flow loss may have been due to a problem with a recovery compressor.

As before the quads come back on line with a local reset

Used the same sequence as before on the right Dipole: depower power supply, depower allen bradley, wait 1 min. repower allen bradley and repower power supply and it worked the first time.

All magnets operational