#### User name stevensm

#### Log entry time 20:46:01 on August15,2005

#### Entry number 152087

#### This entry is a followup to: 152082

keyword=Right Dipole Lead A lead flow trip

I responded to a call when the right dipole failed on lead flow and wouldn't reset. We entered the hall and discovered that the Lead A lead flow controller had forgotten its set point number one, which should have been 40 standard liters per minute. This is the hardware interlock for lead flow. The setpoint value had become corrupted and had a value of about 594 standard liters. I reprogrammed the correct value into the Newport controller and everything reset correctly.