The data logging ozone monitor was displaying an offset (3ppm high). Jennifer and I entered the hall and replaced the monitor. Offset no longer present. Both monitors displayed 0 ppm.