Subject: Q1 power supply comm error

From: "Hall A Electronic Logbook" <adaq-noreturn@jlab.org>

Date: Thu, 12 Apr 2007 13:57:31 -0400

To: folts@jlab.org, segal@jlab.org

#### User name M. Stevens

#### Log entry time 08:06:56 on March 19, 2007

#### Entry number 196564

keyword=TECH Left Q1 trip, odd reset

Left q1 arbitrarily tripped. No disturbances in the cryogenics. Failure mode: Magnet appeared to be communicating; in "off" mode we could change the polarity but we couldn't ramp the magnet; Unresetable "Spare Interlock" and "Max Current Set" flags. Reset everything on the gui and the green button in the counting house center room to no avail. On a long shot we reset IOCHA48, and that worked, but we still see the above flags. These flags are informational only and not among the hardware interlock faults. We ramped the magnet to the needed current and resumed operations.

A copy of this log entry has been emailed to: spiegel@jlab.org, fansler@jlab.org, folts@jlab.org, stevensm@jlab.org