#### User name stevensm

#### Log entry time 11:59:36 on August 7,2005

#### Entry number 151036

keyword=Q1 vacuum interlock

Gary called me about 1100 and said left Q1 wouldn't reset. By phone I had him step through the normal reset procedure, and resetting the Q1 instrumentation rack by cycling UPS power to it, to no avail. I came in to help him. On arrival I observed the magnet tripped on its insulating vacuum interlock. The vacuum meter display was flickering randomly, with no discernable values. I cycled power to the vacuum meter (it is not plugged into the UPS). The meter powered up normally, and displayed a good vacuum value (0). The interlock then cleared, and the magnet ramped normally. We looked at plugging the vacuum meter into the UPS, but the UPS is already chock-a-block with everything including the ice management equipment plugged into it. I will discuss this situation with the tech staff, and change it somehow. We need a simple way to cycle power to all of the interlocked meters, such as with the UPS switch.