#### User name Fansler

#### Log entry time 13:49:01 on August 09, 2006

#### Entry number 178646

keyword=RQ1 N2 Level

RQ1 N2 level has been decreasing gradually for several days. In controlled access, we observed heavy frost on the N2 return lines. The liquid level meter read zero, but heavy frost and ice indicated the magnet was overfilling. I cycled power on the liquid level meter. Upon re-power the meter read 60% and slowly floated back to zero. I replaced the meter and calibrated in place. The indications seem to be steady. Limited time during the access prevented a wet calibration of the probe. The in place calibration is accurate +/- 10%.