Subject: Q1 power supply comm error

From: "Hall A Electronic Logbook" <adaq-noreturn@jlab.org>

Date: Thu, 12 Apr 2007 13:57:31 -0400

To: folts@jlab.org, segal@jlab.org

#### User name Fansler/Stevens

#### Log entry time 11:45:30 on April 23, 2007

#### Entry number 199754

keyword=R Q1 Lead Flow Offset

We noticed an offset in the lead flow readbacks on right Q1. The local displays were were 20% lower than the the readbacks on the tools page. Mark narrowed the fault to the isolation modules at the output of the Lead Flow Controllers (Newport process Meters). We replaced the isolation modules and corrected the offset.

A copy of this log entry has been emailed to: fansler@jlab.org, stevensm@jlab.org