Right Q1 He flow A readback was noisy. Actual flow was steady. Replaced newport controller on Q1 right He flow A. Flow appears to be steady.