#### User name folts

#### Log entry time 14:01:38 on November 14, 2005

#### Entry number 160169

keyword=tech/Rt Q1 LN2 LL

Rt Q1 LN2 LL was found to be 0

Replaced meter, set probe length to 18.5 and set min for cal when res. was known to be empty.

Force cooled shield with manual bypass valve, shield is cooling

A copy of this log entry has been emailed to: fansler@jlab.org