#### User name dezern

#### Log entry time 15:26:13 on April 27, 2007

#### Entry number 200292

keyword=lead flow-magnet quench

Lost lead flow to hall. All magnets tripped. We reset all magnets without issue except right Q1. Right Q1 was displaying a comm error. We tried cycling power and resetting the Saclay rack with no joy. We removed and reinstalled the RS-232 plug at the back of control panel and cycled power on the Saclay rack UPS, this cleared the fault.

A copy of this log entry has been emailed to: fansler@jlab.org, dezern@jlab.org, stevensm@jlab.org