Left Q3 Heat Load				1-20-07

The inlet temperature of the left Q3 was increasing to 60k within 3 minutes of the valve closing.  Further investigation revealed an ice ball on the y intersection of the return u-tube.  Pumping the u-tube did not appear to reduce the ice.  We pulled the u-tube and sent it to cryo for repair.  Cryo discovered a short created from overheating during construction.  The u-tube was repaired and leak checked.  It will be re-installed after the February experiment.