Right Q3 Heat Load				1-20-07

We discover a heat load in the right Q3.  We pumped the u-tube and discover a vacuum of 235 mtorr.  We pumped the u-tube to 0 mtorr, and removed the pump.  Four hours later the vacuum was back up to over 200 mtorr.  We leak checked the u-tube, and found no leaks.  After several pumping iterations and leak checks we removed the u-tube.  We suspected the CVI pump out seal, upon inspection we discover trash on the o-ring.  We replaced the pump out and o-ring.  We pumped and purged the u-tube, and leaked checked it.  It tested ok, no leaks.  The u-tube will be re-installed after the February experiment.