We cannot get the left spectrometer GUI to go above 3.16 GeV/c.
The problem is in the software and we are now confirming what the correct software limit should be. Pam at MCC will set the correct number.

Left HRS limit has been reset to 4.3 GeV/c. The 3.15 number I was stuck at is the right limit. We do not know what the left limit was
set to the same number, but it has happened at least once before.

1) To set the left arm to high momentum (4.15 GeV/c) you first type
in this momentum to get the quads correct, but you will need to adjust
the dipole current by hand. Presently, in order to get the 1591 Amps
that are needed you need to set the current to 1701.5 Amp by hand. This number may drift slightly. We everything is correct, all momentum indicators will be green.

2) The IOC for the magnets has been loosing its network connection.
To fix it go to http://hareboot10/ (see passwords list in counting house) and do an immediate reboot. This will take 5-10 min. If you want to know if the IOC is coming back type "ping iocha14" in an xterm. After a reboot you will need to turn the quads back on and probably need to repeat item 1.